## Instructions
Sandia Primary Standards Lab - Excel CurveFit Uncertainty Calculator
v0.1
uncertainty@sandia.gov
This spreadsheet calculates uncertaint in curve fit applications with arbitrary fit models, and accounts for known uncertainty in y. Currently implemented fit models are listed below. Other arbitrary models may be added by defining LAMBDAs to compute the model and its Jacobian (See Formula Reference tab for details). No Excel Macros are used, however the calculations make use of Excel's LAMBDA function which is only available in Office 365 or 2024.
Usage
Features | • Input known uncertainty in y, or leave blank to use average residual as y uncertainty • Compute uncertainty in fit parameters • Plot confidence and prediction bands • Compute predicted y and its uncertainty at any x values
Available Fit Models | • Line • Quadratic • Cubic • Exponential Decay • Damped Sine
Instructions | 1) Enter measured x, y, and optionally u(y) values in the table 2) Select a fit model from the dropdown 3) If a solution did not compute properly, try adjusting the initial guess values 4) In the predictions table, enter x values at which to predict y and its uncertainty
Results | • Fit parameters: These are the fit model coefficients that minimize the sum of squared residuals • Fit parameter uncertainties: Standard uncertainty in the fit parameters • Covariance Matrix: The covariance of the fit parameters • Confidence Band: The expanded uncertainty of the location of the fit line at the given x value • Prediction Band: Confidence band root-sum-squared with the uncertainty in a new measurement. Measured values should fall within the prediction band with the provided level of confidence.
Worksheets
Curve Fit Uncertainty | The calculator
Examples | Some example data to try and used for validating the calculations
Formula Reference | Lists the LAMBDA functions used in Excel's Name Manager for readability. Hidden by default.
Lookup | Provides allowable options for dropdown menus. Hidden by default.
About
Sandia National Laboratories is a multimission laboratory managed and operated by National Technology & Engineering Solutions of Sandia, LLC, a wholly owned subsidiary of Honeywell International Inc., for the U.S. Department of Energy’s National Nuclear Security Administration under contract DE-NA0003525.
SAND2026-16029O

## Curve Fit Uncertainty
1) Enter measured x, y, and optionally u(y) values |  |  |  |  | 2) Select fit model and provide initial guess |  |  |  |  | 3) Enter X values at which to predict Y and its expanded uncertainty | [11 empty] | ← Plot data in hidden columns here
x | y | uy | Predicted | Residual | Fit Model |  |  |  |  | X | Predicted Y | Confidence Band ± | Prediction Band ± | Plot Fit X | Plot Fit Y | ±Confidence Band (k=1) | ±Prediction Band (k=1) | Plot Fit Conf- | Plot Fit Conf+ | Plot Pred- | Plot Pred+
0 | 4.7662 |  | 4.7934 | 0.027155 | Fit Model | Exponential | y= | a*exp(-x/b) |  | 0 | 4.7934 | 0.19251 | 0.28321 | 0 | 4.7934 | 0.085098 | 0.12519 | 4.6008 | 4.9859 | 4.5101 | 5.0766
1 | 2.9278 |  | 2.9406 | 0.012763 | Level of Confidence | 95% |  |  |  | 5 | 0.41648 | 0.065369 | 0.21777 | 0.20408 | 4.3384 | 0.070846 | 0.11598 | 4.1782 | 4.4987 | 4.0761 | 4.6008
2 | 1.9706 |  | 1.8039 | -0.16666 | Initial Guess | 1 | 1 | 1 | 1 | 10 | 0.036187 | 0.011911 | 0.20807 | 0.40816 | 3.9267 | 0.060565 | 0.11 | 3.7897 | 4.0637 | 3.6778 | 4.1755
3 | 1.0177 |  | 1.1067 | 0.088955 | [9 empty] | 0.61224 | 3.554 | 0.05377 | 0.10641 | 3.4323 | 3.6756 | 3.3133 | 3.7947
4 | 0.6841 |  | 0.6789 | -0.0052044 | Fit Results | [8 empty] | 0.81633 | 3.2167 | 0.049797 | 0.10446 | 3.104 | 3.3293 | 2.9804 | 3.453
5 | 0.4564 |  | 0.41648 | -0.03992 |  | a | b | c | d |  |  |  |  | 1.0204 | 2.9114 | 0.047853 | 0.10355 | 2.8031 | 3.0196 | 2.6771 | 3.1456
6 | 0.1527 |  | 0.2555 | 0.1028 | Fit Parameters | 4.7934 | 2.0465 | [6 empty] | 1.2245 | 2.6351 | 0.047169 | 0.10323 | 2.5284 | 2.7418 | 2.4015 | 2.8686
7 | 0.0024 |  | 0.15674 | 0.15434 | Standard Uncertainties | 0.085098 | 0.064504 | 0 | 0 |  |  |  |  | 1.4286 | 2.385 | 0.047127 | 0.10321 | 2.2784 | 2.4916 | 2.1515 | 2.6185
8 | 0.0935 |  | 0.096154 | 0.0026538 | Expanded Uncertainties | 0.19251 | 0.14592 | 0 | 0 |  |  |  |  | 1.6327 | 2.1586 | 0.047305 | 0.10329 | 2.0516 | 2.2656 | 1.9249 | 2.3923
9 | 0.0098 |  | 0.058987 | 0.049187 | [9 empty] | 1.8367 | 1.9537 | 0.047447 | 0.10336 | 1.8464 | 2.0611 | 1.7199 | 2.1876
10 | 0.0068 |  | 0.036187 | 0.029387 | Covariance Matrix | 0.0072417 | -0.0028723 | [6 empty] | 2.0408 | 1.7683 | 0.047418 | 0.10335 | 1.661 | 1.8756 | 1.5345 | 2.0021
[6 empty] | -0.0028723 | 0.0041607 | [6 empty] | 2.2449 | 1.6005 | 0.047157 | 0.10323 | 1.4938 | 1.7072 | 1.367 | 1.834
[14 empty] | 2.449 | 1.4486 | 0.046647 | 0.10299 | 1.3431 | 1.5541 | 1.2156 | 1.6816
[14 empty] | 2.6531 | 1.3111 | 0.045898 | 0.10266 | 1.2073 | 1.4149 | 1.0789 | 1.5433
[14 empty] | 2.8571 | 1.1867 | 0.044934 | 0.10223 | 1.085 | 1.2883 | 0.9554 | 1.4179
[5 empty] | # fit parameters | 2 | [7 empty] | 3.0612 | 1.074 | 0.043785 | 0.10173 | 0.97499 | 1.1731 | 0.84391 | 1.3042
[5 empty] | # data points | 11 | [7 empty] | 3.2653 | 0.9721 | 0.042482 | 0.10118 | 0.876 | 1.0682 | 0.74322 | 1.201
[5 empty] | Sum of squared residuals | 0.075888 | [7 empty] | 3.4694 | 0.87984 | 0.041055 | 0.10059 | 0.78697 | 0.97271 | 0.6523 | 1.1074
[5 empty] | Standard Error Syx | 0.091826 | [7 empty] | 3.6735 | 0.79633 | 0.039535 | 0.099975 | 0.7069 | 0.88577 | 0.57018 | 1.0225
[5 empty] | Coverage Factor k | 2.2622 | [7 empty] | 3.8776 | 0.72076 | 0.037946 | 0.099358 | 0.63491 | 0.8066 | 0.49599 | 0.94552
[14 empty] | 4.0816 | 0.65235 | 0.036314 | 0.098746 | 0.5702 | 0.7345 | 0.42897 | 0.87573
[14 empty] | 4.2857 | 0.59043 | 0.034657 | 0.098148 | 0.51204 | 0.66883 | 0.36841 | 0.81246
[14 empty] | 4.4898 | 0.5344 | 0.032993 | 0.097573 | 0.45976 | 0.60903 | 0.31367 | 0.75512
[14 empty] | 4.6939 | 0.48368 | 0.031338 | 0.097026 | 0.41279 | 0.55457 | 0.26419 | 0.70317
[14 empty] | 4.898 | 0.43777 | 0.029703 | 0.096511 | 0.37058 | 0.50497 | 0.21945 | 0.65609
[14 empty] | 5.102 | 0.39622 | 0.028099 | 0.096029 | 0.33266 | 0.45979 | 0.17899 | 0.61346
[14 empty] | 5.3061 | 0.35862 | 0.026535 | 0.095583 | 0.29859 | 0.41864 | 0.14239 | 0.57484
[14 empty] | 5.5102 | 0.32458 | 0.025016 | 0.095172 | 0.26799 | 0.38117 | 0.10929 | 0.53988
[14 empty] | 5.7143 | 0.29378 | 0.023548 | 0.094797 | 0.24051 | 0.34704 | 0.07933 | 0.50822
[14 empty] | 5.9184 | 0.26589 | 0.022134 | 0.094456 | 0.21582 | 0.31596 | 0.052219 | 0.47957
[14 empty] | 6.1224 | 0.24066 | 0.020778 | 0.094147 | 0.19366 | 0.28766 | 0.027682 | 0.45363
[14 empty] | 6.3265 | 0.21782 | 0.01948 | 0.093869 | 0.17375 | 0.26188 | 0.0054696 | 0.43016
[14 empty] | 6.5306 | 0.19714 | 0.018242 | 0.09362 | 0.15588 | 0.23841 | -0.01464 | 0.40893
[14 empty] | 6.7347 | 0.17843 | 0.017064 | 0.093398 | 0.13983 | 0.21704 | -0.032848 | 0.38971
[14 empty] | 6.9388 | 0.1615 | 0.015946 | 0.0932 | 0.12543 | 0.19757 | -0.049335 | 0.37233
[14 empty] | 7.1429 | 0.14617 | 0.014887 | 0.093025 | 0.11249 | 0.17985 | -0.064266 | 0.35661
[14 empty] | 7.3469 | 0.1323 | 0.013885 | 0.09287 | 0.10089 | 0.16371 | -0.077789 | 0.34238
[14 empty] | 7.551 | 0.11974 | 0.01294 | 0.092733 | 0.090469 | 0.14901 | -0.090036 | 0.32952
[14 empty] | 7.7551 | 0.10838 | 0.012049 | 0.092613 | 0.081119 | 0.13563 | -0.10113 | 0.31788
[14 empty] | 7.9592 | 0.098091 | 0.011211 | 0.092508 | 0.072729 | 0.12345 | -0.11118 | 0.30736
[14 empty] | 8.1633 | 0.088781 | 0.010424 | 0.092416 | 0.065201 | 0.11236 | -0.12028 | 0.29784
[14 empty] | 8.3673 | 0.080355 | 0.009685 | 0.092335 | 0.058446 | 0.10226 | -0.12852 | 0.28923
[14 empty] | 8.5714 | 0.072728 | 0.0089926 | 0.092265 | 0.052386 | 0.093071 | -0.13599 | 0.28145
[14 empty] | 8.7755 | 0.065826 | 0.0083445 | 0.092204 | 0.046949 | 0.084702 | -0.14275 | 0.27441
[14 empty] | 8.9796 | 0.059578 | 0.0077385 | 0.092151 | 0.042073 | 0.077084 | -0.14888 | 0.26804
[14 empty] | 9.1837 | 0.053924 | 0.0071724 | 0.092106 | 0.037699 | 0.070149 | -0.15443 | 0.26228
[14 empty] | 9.3878 | 0.048806 | 0.0066441 | 0.092066 | 0.033776 | 0.063836 | -0.15946 | 0.25707
[14 empty] | 9.5918 | 0.044174 | 0.0061515 | 0.092032 | 0.030258 | 0.058089 | -0.16402 | 0.25236
[14 empty] | 9.7959 | 0.039981 | 0.0056926 | 0.092002 | 0.027104 | 0.052859 | -0.16814 | 0.2481
[5 empty] | Use the plot's filter button to show the confidence band and predicted points. | [8 empty] | 10 | 0.036187 | 0.0052654 | 0.091977 | 0.024275 | 0.048098 | -0.17188 | 0.24425

## Examples
RLC Circuit Response |  |  |  | Temperature coefficient of resistance |  |  |  |  | RC Circuit Response
Model: | Damped Sine |  |  | Model: | Linear, Quadratic, or Cubic |  |  |  | Model: | Exponential
Initial Guess | 2 | 2e-07 | 8000000 | Initial Guess | 1 | 1 | 1 | 1
x | y |  |  | x | y |  |  |  | x | y
0 | -0.011364 |  |  | 16.1899 | 200.0417 |  |  |  | 0 | 4.7662
4.008e-09 | 0.080924 |  |  | 17.2607 | 200.0433 |  |  |  | 1 | 2.9278
8.016e-09 | 0.13939 |  |  | 18.3324 | 200.0454 |  |  |  | 2 | 1.9706
1.2024e-08 | 0.19488 |  |  | 19.3391 | 200.0478 |  |  |  | 3 | 1.0177
1.6032e-08 | 0.29534 |  |  | 20.4266 | 200.0491 |  |  |  | 4 | 0.6841
2.004e-08 | 0.31699 |  |  | 21.4117 | 200.0513 |  |  |  | 5 | 0.4564
2.4048e-08 | 0.35727 |  |  | 22.4135 | 200.0536 |  |  |  | 6 | 0.1527
2.8056e-08 | 0.41582 |  |  | 23.4931 | 200.0548 |  |  |  | 7 | 0.0024
3.2064e-08 | 0.48013 |  |  | 24.5274 | 200.0565 |  |  |  | 8 | 0.0935
3.6072e-08 | 0.50545 |  |  | 25.5361 | 200.058 |  |  |  | 9 | 0.0098
4.008e-08 | 0.61021 |  |  | 26.5547 | 200.0594 |  |  |  | 10 | 0.0068
4.4088e-08 | 0.58452 |  |  | 27.5919 | 200.0605
4.8096e-08 | 0.61018 |  |  | 28.6312 | 200.0616
5.2104e-08 | 0.64698 |  |  | 29.698 | 200.0626
5.6112e-08 | 0.6692
6.012e-08 | 0.76226
6.4128e-08 | 0.76258
6.8136e-08 | 0.79086
7.2144e-08 | 0.84345
7.6152e-08 | 0.82952
8.016e-08 | 0.88899
8.4168e-08 | 0.88735
8.8176e-08 | 0.86623
9.2184e-08 | 0.88056
9.6192e-08 | 0.95014
1.002e-07 | 0.94391
1.0421e-07 | 0.94549
1.0822e-07 | 0.95914
1.1222e-07 | 0.8996
1.1623e-07 | 0.90918
1.2024e-07 | 0.89663
1.2425e-07 | 0.93639
1.2826e-07 | 0.9211
1.3226e-07 | 0.94419
1.3627e-07 | 0.89048
1.4028e-07 | 0.89942
1.4429e-07 | 0.91685
1.483e-07 | 0.90317
1.523e-07 | 0.87635
1.5631e-07 | 0.954
1.6032e-07 | 0.87305
1.6433e-07 | 0.86154
1.6834e-07 | 0.83808
1.7234e-07 | 0.83661
1.7635e-07 | 0.89779
1.8036e-07 | 0.77362
1.8437e-07 | 0.79636
1.8838e-07 | 0.74917
1.9238e-07 | 0.75337
1.9639e-07 | 0.74862
2.004e-07 | 0.742
2.0441e-07 | 0.68631
2.0842e-07 | 0.6965
2.1242e-07 | 0.6814
2.1643e-07 | 0.68888
2.2044e-07 | 0.6085
2.2445e-07 | 0.6795
2.2846e-07 | 0.54626
2.3246e-07 | 0.56495
2.3647e-07 | 0.55687
2.4048e-07 | 0.48261
2.4449e-07 | 0.45386
2.485e-07 | 0.54367
2.5251e-07 | 0.47127
2.5651e-07 | 0.46749
2.6052e-07 | 0.33997
2.6453e-07 | 0.38291
2.6854e-07 | 0.3747
2.7255e-07 | 0.31589
2.7655e-07 | 0.3519
2.8056e-07 | 0.33638
2.8457e-07 | 0.3692
2.8858e-07 | 0.33112
2.9259e-07 | 0.26176
2.9659e-07 | 0.27142
3.006e-07 | 0.27598
3.0461e-07 | 0.22953
3.0862e-07 | 0.23453
3.1263e-07 | 0.15549
3.1663e-07 | 0.1502
3.2064e-07 | 0.16558
3.2465e-07 | 0.096707
3.2866e-07 | 0.15435
3.3267e-07 | 0.10716
3.3667e-07 | 0.033915
3.4068e-07 | 0.056975
3.4469e-07 | 0.060335
3.487e-07 | 0.032186
3.5271e-07 | 0.068424
3.5671e-07 | 0.031988
3.6072e-07 | 0.041189
3.6473e-07 | -0.0027879
3.6874e-07 | -0.056244
3.7275e-07 | -0.050763
3.7675e-07 | -0.053974
3.8076e-07 | -0.04501
3.8477e-07 | -0.033029
3.8878e-07 | -0.056478
3.9279e-07 | -0.098738
3.9679e-07 | -0.09383
4.008e-07 | -0.10613
4.0481e-07 | -0.11512
4.0882e-07 | -0.12281
4.1283e-07 | -0.11459
4.1683e-07 | -0.13814
4.2084e-07 | -0.10366
4.2485e-07 | -0.1077
4.2886e-07 | -0.15337
4.3287e-07 | -0.13484
4.3687e-07 | -0.15121
4.4088e-07 | -0.14556
4.4489e-07 | -0.14186
4.489e-07 | -0.15057
4.5291e-07 | -0.082824
4.5691e-07 | -0.1778
4.6092e-07 | -0.14318
4.6493e-07 | -0.19766
4.6894e-07 | -0.13206
4.7295e-07 | -0.1372
4.7695e-07 | -0.071458
4.8096e-07 | -0.20289
4.8497e-07 | -0.16585
4.8898e-07 | -0.1556
4.9299e-07 | -0.20344
4.9699e-07 | -0.14794
5.01e-07 | -0.2027
5.0501e-07 | -0.13135
5.0902e-07 | -0.22882
5.1303e-07 | -0.21222
5.1703e-07 | -0.1378
5.2104e-07 | -0.095758
5.2505e-07 | -0.14806
5.2906e-07 | -0.18532
5.3307e-07 | -0.14703
5.3707e-07 | -0.19539
5.4108e-07 | -0.11737
5.4509e-07 | -0.071994
5.491e-07 | -0.15641
5.5311e-07 | -0.10965
5.5711e-07 | -0.089519
5.6112e-07 | -0.058503
5.6513e-07 | -0.099744
5.6914e-07 | -0.11608
5.7315e-07 | -0.092336
5.7715e-07 | -0.079843
5.8116e-07 | -0.1243
5.8517e-07 | -0.11112
5.8918e-07 | -0.097008
5.9319e-07 | -0.090863
5.9719e-07 | -0.11812
6.012e-07 | -0.11325
6.0521e-07 | -0.080642
6.0922e-07 | -0.1182
6.1323e-07 | -0.12845
6.1723e-07 | -0.093834
6.2124e-07 | -0.095012
6.2525e-07 | -0.14882
6.2926e-07 | -0.075347
6.3327e-07 | -0.065415
6.3727e-07 | -0.043125
6.4128e-07 | -0.013416
6.4529e-07 | -0.073573
6.493e-07 | -0.027793
6.5331e-07 | -0.039315
6.5731e-07 | -0.049439
6.6132e-07 | -0.093915
6.6533e-07 | -0.007409
6.6934e-07 | -0.085761
6.7335e-07 | -0.04344
6.7735e-07 | -0.094017
6.8136e-07 | 0.058646
6.8537e-07 | 0.012066
6.8938e-07 | -0.031699
6.9339e-07 | -0.08554
6.9739e-07 | -0.037719
7.014e-07 | 0.0075115
7.0541e-07 | 0.044816
7.0942e-07 | -0.019106
7.1343e-07 | -0.013986
7.1743e-07 | 0.0037521
7.2144e-07 | -0.038882
7.2545e-07 | 0.023124
7.2946e-07 | 0.0091938
7.3347e-07 | 0.0076298
7.3747e-07 | 0.0016722
7.4148e-07 | -0.054773
7.4549e-07 | -0.0047452
7.495e-07 | 0.037397
7.5351e-07 | 0.021618
7.5752e-07 | -0.033739
7.6152e-07 | 0.049524
7.6553e-07 | -0.026318
7.6954e-07 | -0.0033987
7.7355e-07 | -0.022617
7.7756e-07 | 0.0059893
7.8156e-07 | 0.0092967
7.8557e-07 | 0.057685
7.8958e-07 | -0.017614
7.9359e-07 | 0.052896
7.976e-07 | 0.049258
8.016e-07 | 0.061476
8.0561e-07 | -0.015827
8.0962e-07 | 0.010143
8.1363e-07 | 0.036089
8.1764e-07 | 0.022044
8.2164e-07 | -0.021745
8.2565e-07 | 0.020472
8.2966e-07 | 0.017557
8.3367e-07 | 0.054799
8.3768e-07 | 0.023108
8.4168e-07 | -0.015181
8.4569e-07 | 0.026305
8.497e-07 | 0.071243
8.5371e-07 | 0.072789
8.5772e-07 | 0.012332
8.6172e-07 | 0.0073903
8.6573e-07 | 0.042362
8.6974e-07 | -0.0021251
8.7375e-07 | -0.016742
8.7776e-07 | 0.031788
8.8176e-07 | 0.068765
8.8577e-07 | 0.063994
8.8978e-07 | 0.038496
8.9379e-07 | 0.034778
8.978e-07 | 0.011155
9.018e-07 | 0.059816
9.0581e-07 | 0.003226
9.0982e-07 | -0.0092947
9.1383e-07 | -0.00051857
9.1784e-07 | 0.024138
9.2184e-07 | -0.026367
9.2585e-07 | -0.0058967
9.2986e-07 | -0.015059
9.3387e-07 | 0.00093877
9.3788e-07 | -0.0027446
9.4188e-07 | 0.045848
9.4589e-07 | -0.0035304
9.499e-07 | 0.0070842
9.5391e-07 | -0.034421
9.5792e-07 | -0.015716
9.6192e-07 | -0.026869
9.6593e-07 | 0.033894
9.6994e-07 | 0.022294
9.7395e-07 | 0.0022344
9.7796e-07 | 0.049418
9.8196e-07 | 0.023771
9.8597e-07 | -0.036111
9.8998e-07 | 0.026
9.9399e-07 | 0.04122
9.98e-07 | 0.01008
1.002e-06 | -0.00030574
1.006e-06 | 0.055026
1.01e-06 | -0.014994
1.014e-06 | 0.03601
1.018e-06 | 0.016892
1.022e-06 | -0.087092
1.0261e-06 | -0.015456
1.0301e-06 | -0.012982
1.0341e-06 | -0.02451
1.0381e-06 | 0.063131
1.0421e-06 | -0.06891
1.0461e-06 | 0.068262
1.0501e-06 | 0.01014
1.0541e-06 | 0.024915
1.0581e-06 | -0.040029
1.0621e-06 | 0.034039
1.0661e-06 | -0.050477
1.0701e-06 | -0.0095444
1.0741e-06 | 0.060629
1.0782e-06 | 0.00090911
1.0822e-06 | -0.015487
1.0862e-06 | -0.022171
1.0902e-06 | 0.0082957
1.0942e-06 | -0.024716
1.0982e-06 | 0.017907
1.1022e-06 | -0.0027473
1.1062e-06 | -0.021962
1.1102e-06 | -0.014855
1.1142e-06 | -0.0047595
1.1182e-06 | -0.0042634
1.1222e-06 | 0.028545
1.1263e-06 | 0.018967
1.1303e-06 | 0.0036948
1.1343e-06 | -0.027905
1.1383e-06 | -0.018399
1.1423e-06 | 0.034793
1.1463e-06 | -0.0039387
1.1503e-06 | -0.007758
1.1543e-06 | 0.0086313
1.1583e-06 | 0.044418
1.1623e-06 | 0.00234
1.1663e-06 | -0.018967
1.1703e-06 | 0.069602
1.1743e-06 | -0.0094511
1.1784e-06 | 0.0027891
1.1824e-06 | -0.012342
1.1864e-06 | -0.016261
1.1904e-06 | 0.024562
1.1944e-06 | 0.016235
1.1984e-06 | -0.027619
1.2024e-06 | 0.039738
1.2064e-06 | -0.0031789
1.2104e-06 | -0.0059623
1.2144e-06 | 0.011764
1.2184e-06 | -0.032242
1.2224e-06 | 0.028026
1.2265e-06 | -0.017829
1.2305e-06 | -0.016805
1.2345e-06 | 0.012988
1.2385e-06 | -0.021654
1.2425e-06 | 0.0098504
1.2465e-06 | 0.043738
1.2505e-06 | -0.011727
1.2545e-06 | -0.034755
1.2585e-06 | 0.041279
1.2625e-06 | -0.034848
1.2665e-06 | -0.026697
1.2705e-06 | 0.013977
1.2745e-06 | -0.021079
1.2786e-06 | -0.026555
1.2826e-06 | 0.064226
1.2866e-06 | -0.11491
1.2906e-06 | -0.015752
1.2946e-06 | 0.029518
1.2986e-06 | -0.032317
1.3026e-06 | -0.01251
1.3066e-06 | -0.015884
1.3106e-06 | -0.042104
1.3146e-06 | -0.016946
1.3186e-06 | -0.0076881
1.3226e-06 | -0.046238
1.3267e-06 | -0.036427
1.3307e-06 | 0.020783
1.3347e-06 | 0.013704
1.3387e-06 | -0.039952
1.3427e-06 | 0.0037546
1.3467e-06 | 0.03681
1.3507e-06 | 0.026862
1.3547e-06 | 0.0034002
1.3587e-06 | -0.03269
1.3627e-06 | 0.028436
1.3667e-06 | 0.006384
1.3707e-06 | 0.0074985
1.3747e-06 | -0.018067
1.3788e-06 | 0.013248
1.3828e-06 | -9.1237e-05
1.3868e-06 | 0.026773
1.3908e-06 | 0.039023
1.3948e-06 | -0.040454
1.3988e-06 | -0.024455
1.4028e-06 | -0.017509
1.4068e-06 | -0.05201
1.4108e-06 | 0.027953
1.4148e-06 | -0.026219
1.4188e-06 | -0.027324
1.4228e-06 | 0.04977
1.4269e-06 | 0.033826
1.4309e-06 | 0.011824
1.4349e-06 | -0.024701
1.4389e-06 | 0.032558
1.4429e-06 | 0.0072844
1.4469e-06 | 0.0084574
1.4509e-06 | 0.028529
1.4549e-06 | 0.0077814
1.4589e-06 | 0.00052637
1.4629e-06 | 0.028618
1.4669e-06 | -0.031867
1.4709e-06 | -0.030194
1.4749e-06 | 0.013255
1.479e-06 | -0.0051098
1.483e-06 | 0.0049096
1.487e-06 | 0.03059
1.491e-06 | -0.057476
1.495e-06 | 0.069936
1.499e-06 | -0.033556
1.503e-06 | -0.018345
1.507e-06 | 0.0049928
1.511e-06 | 0.012169
1.515e-06 | -0.032853
1.519e-06 | 0.01669
1.523e-06 | 0.01649
1.5271e-06 | 0.037812
1.5311e-06 | 0.0055879
1.5351e-06 | -0.0084662
1.5391e-06 | -0.017185
1.5431e-06 | -0.017418
1.5471e-06 | -0.027519
1.5511e-06 | -0.00096448
1.5551e-06 | 0.039574
1.5591e-06 | -0.055743
1.5631e-06 | 0.050409
1.5671e-06 | -0.0042286
1.5711e-06 | -0.050089
1.5752e-06 | 0.056445
1.5792e-06 | -0.019861
1.5832e-06 | -0.022651
1.5872e-06 | -0.052532
1.5912e-06 | 0.027852
1.5952e-06 | 0.039008
1.5992e-06 | 0.022312
1.6032e-06 | 0.0051734
1.6072e-06 | 0.030532
1.6112e-06 | 0.00021364
1.6152e-06 | 0.034554
1.6192e-06 | -0.0048784
1.6232e-06 | 0.02574
1.6273e-06 | -0.030148
1.6313e-06 | -0.021155
1.6353e-06 | 0.029669
1.6393e-06 | -0.080526
1.6433e-06 | 0.033083
1.6473e-06 | -0.0093927
1.6513e-06 | -0.018377
1.6553e-06 | -0.0185
1.6593e-06 | 0.0067022
1.6633e-06 | 0.031069
1.6673e-06 | 0.051046
1.6713e-06 | -0.030618
1.6754e-06 | -0.035445
1.6794e-06 | 0.025837
1.6834e-06 | 0.0099381
1.6874e-06 | -0.005839
1.6914e-06 | 0.0064227
1.6954e-06 | 0.010486
1.6994e-06 | -0.0067994
1.7034e-06 | -0.0014005
1.7074e-06 | -0.00724
1.7114e-06 | 0.045304
1.7154e-06 | 0.016228
1.7194e-06 | 0.015164
1.7234e-06 | -0.046597
1.7275e-06 | -0.027809
1.7315e-06 | -0.0078295
1.7355e-06 | 0.035595
1.7395e-06 | -0.025121
1.7435e-06 | 0.059821
1.7475e-06 | 0.045615
1.7515e-06 | 0.0026531
1.7555e-06 | -0.051168
1.7595e-06 | 0.016312
1.7635e-06 | -0.0098824
1.7675e-06 | 0.034687
1.7715e-06 | 0.042639
1.7756e-06 | 0.010661
1.7796e-06 | 0.0037917
1.7836e-06 | 0.035066
1.7876e-06 | -0.022242
1.7916e-06 | 0.0075858
1.7956e-06 | 0.0093161
1.7996e-06 | -0.038127
1.8036e-06 | -0.02842
1.8076e-06 | 0.013366
1.8116e-06 | -0.012062
1.8156e-06 | 0.017835
1.8196e-06 | 0.064831
1.8236e-06 | -0.0081377
1.8277e-06 | -0.014038
1.8317e-06 | 0.0078365
1.8357e-06 | -0.048299
1.8397e-06 | -0.053324
1.8437e-06 | -0.032503
1.8477e-06 | 0.00085579
1.8517e-06 | -0.031453
1.8557e-06 | -0.0028625
1.8597e-06 | -0.016236
1.8637e-06 | -0.056755
1.8677e-06 | -0.0099962
1.8717e-06 | -0.021214
1.8758e-06 | -0.035595
1.8798e-06 | -0.051729
1.8838e-06 | 0.0045573
1.8878e-06 | 0.0078895
1.8918e-06 | 0.0016694
1.8958e-06 | 0.020895
1.8998e-06 | 0.0568
1.9038e-06 | -0.05633
1.9078e-06 | -0.024021
1.9118e-06 | -0.029938
1.9158e-06 | 0.028035
1.9198e-06 | 0.0058034
1.9238e-06 | -0.033938
1.9279e-06 | -0.062061
1.9319e-06 | -0.021193
1.9359e-06 | -0.013255
1.9399e-06 | 0.077962
1.9439e-06 | 0.0321
1.9479e-06 | 0.0009399
1.9519e-06 | -0.0075666
1.9559e-06 | -0.024391
1.9599e-06 | -0.040338
1.9639e-06 | 0.041423
1.9679e-06 | 0.046385
1.9719e-06 | 0.011738
1.976e-06 | 0.0056472
1.98e-06 | -0.06846
1.984e-06 | 0.020605
1.988e-06 | 0.028002
1.992e-06 | 0.021992
1.996e-06 | 0.0023709
2e-06 | 0.017398

## Lookup (hidden)
Line
Quadratic
Cubic
Exponential
Damped Sine

## Formula Reference (hidden)
Formula Reference - These are the same formulas as defined in Name Manager, duplicated here for readability and easier editing.
Curve Fitting Functions.
Curve fit is calculated using Gauss-Newton Method, which requires a function for calculating the fit model and for calculating its Jacobian at x.
| DEL | LAMBDA(x,y,uy,p,func,func_jac,LET(j, func_jac(x, p), jt, TRANSPOSE(j), r, y-func(x, p), w, IF(SUM(uy)>0, DIAG(ARRAYFORMULA(1/uy^2)), DIAG(ONES(uy,1))), minv, MINVERSE(MMULT(MMULT(jt, w), j)), delta1, TRANSPOSE(MMULT(MMULT(jt, w), r)), MMULT(delta1, minv)) ) | Compute delta used in Gauss-Newton method. - Parameters x and y are the measured datapoints. - uy is optional uncertainty in each y. If omitted, the weights are all 1. - p is the array of fit coefficients (e.g. {slope, intercept}) - func is a LAMBDA that calculates y given x and p - func_jac is a LAMBDA that calculates Jacobian of the model given x and p
| GAUSSNEWTON | LAMBDA(x,y,uy,func,func_jac,guess, LET(iterations, 20, initial, TRANSPOSE(guess), loop, REDUCE(initial, SEQUENCE(iterations), LAMBDA(acc, i, LET( coeffs, INDEX(acc, ROWS(acc)-1), b, DEL(x, y, uy, TRANSPOSE(coeffs), func, func_jac), output, VSTACK(acc, ARRAYFORMULA(coeffs-b)), output ) )), INDEX(loop, ROWS(loop)-1) )) | Gauss-Newton Minimization used to minimize sum of squared residuals in curve fit. - x and y are the measured datapoints. - uy is optional uncertainty in each y. If omitted, the weights are all 1. - func is a LAMBDA that calculates y given x and p - func_jac is a LAMBDA that calculates Jacobian of the model given x and p - guess is an intial guess of the fit parameters p
| FIT_COVARIANCE | LAMBDA(x,y,uy,p,resid,func,func_jac, LET(j, func_jac(x, p), jt, TRANSPOSE(j), IF( SUM(uy)=0, LET( c, MINVERSE(MMULT(jt, j)), n, ROWS(x), m, ROWS(coeff), mse, SUM((resid^2)) / (n-m), (c*mse) ), LET( w, MINVERSE(DIAG((1/uy^2))), MMULT(MMULT(jt, w), j) ) ) )) | Compute covariance matrix of curve fit - x and y are the measured datapoints. - uy is optional uncertainty in each y. If omitted, the average residual is used. - p are the fit coefficients found using GAUSSNEWTON - resid is an array of residuals calculated using the fit model - func is a LAMBDA that calculates y given x and p - func_jac is a LAMBDA that calculates Jacobian of the model given x and p
| UCONF | LAMBDA(x,p,cov,func_jac, LET(x, NONBLANK(x), j, func_jac(x, p), n, COUNTA(p), cov, TAKE(cov, n, n), BYROW(j, LAMBDA(ji, LET(delf, TRANSPOSE(ji), mm, MMULT(MMULT(TRANSPOSE(delf), cov), delf), SQRT(mm) ) ) ) ) ) | Calculate confidence band at x - x are the x-values at which to calculate confidence in y - p are the fit coefficients found using GAUSSNEWTON - cov is the covariance matrix found using FIT_COVARIANCE - func_jac is a LAMBDA that calculates Jacobian of the model given x and p
| NONBLANK | LAMBDA(array, FILTER(array, array<>"") ) | Utility function to filter out non-empty values in an array
Fit Models
To add a new fit model, implement the fit model itself in a new MODEL_XXX, and implement its Jacobian in MODEL_XXX_JAC. Then add thse to the lookup functions in Model Lookup Functions defined below. The model and Jacobian LAMBDAs take parameters x and p - where x is the x-value at which to compute y, and p is an array of coefficients (such as slope and intercept) used by the model.
| MODEL_LINE | LAMBDA(x,p, x*INDEX(p,1) + INDEX(p,2)) | Line model: y = a*x + b
| MODEL_LINE_JAC | LAMBDA(x,p, HSTACK( MAP(x, LAMBDA(k, -k)), MAP(x, LAMBDA(k, -1)) ) ) | Jacobian of line model
| MODEL_QUAD | LAMBDA(x,p, x^2*INDEX(p,1) + x*INDEX(p,2) + INDEX(p,3) ) | Quadratic fit model: y = a*x^2 + b*x + c
| MODEL_QUAD_JAC | LAMBDA(x,p, HSTACK( MAP(x, LAMBDA(k, -(k^2))), MAP(x, LAMBDA(k, -k)), MAP(x, LAMBDA(k, -1)) ) ) | Jacobian of quadratic model
| MODEL_CUBIC | LAMBDA(x,p, x^3*INDEX(p,1) + x^2*INDEX(p,2) + x*INDEX(p,3) + INDEX(p,4) ) | Cubic fit model: y = a*x^3 + b*x^2 + c*x + d
| MODEL_CUBIC_JAC | LAMBDA(x,p, HSTACK( MAP(x, LAMBDA(k, -(k^3))), MAP(x, LAMBDA(k, -(k^2))), MAP(x, LAMBDA(k, -k)), MAP(x, LAMBDA(k, -1)) ) ) | Jacobian of cubic model
| MODEL_EXP | LAMBDA(x, p, INDEX(p,1) * EXP(-x/INDEX(p,2)) ) | Exponential decay fit model: y = a*exp(-x/b)
| MODEL_EXP_JAC | LAMBDA(x, p, LET( b, INDEX(p,2), a, INDEX(p,1), HSTACK( MAP(x, LAMBDA(k, -EXP(-k/b))), MAP(x, LAMBDA(k, -a*k/b^2*EXP(-k/b))) ) ) ) | Jacobian of exponential decay model
| MODEL_DSINE | LAMBDA(x, p, LET( a, INDEX(p,1), b, INDEX(p,2), c, INDEX(p,3), a*EXP(-x/b)*SIN(c*x) ) | Damped Sine fit model: y = a*exp(-x/b)*sin(c*x)
| MODEL_DSINE_JAC | LAMBDA(x,p, LET( a, INDEX(p,1), b, INDEX(p,2), c, INDEX(p,3), HSTACK( MAP(x, LAMBDA(k, -EXP(-k/b)*SIN(c*k))), MAP(x, LAMBDA(k, -a*k/b^2*(EXP(-k/b)*SIN(c*k)))), MAP(x, LAMBDA(k, -a*k*EXP(-k/b)*COS(c*k))) ) ) ) | Jacobian of damped sine model
Model Lookup Functions
| MODELNPARAMS | LAMBDA(name, IFS(name="Line", 2, name="Quadratic", 3, name="Cubic", 4, name="Exponential", 2, name="Damped Sine", 3, TRUE, 2 ) ) | Get number of fit parameters for the model
| MODEL | LAMBDA(name, IF(name="Line", MODEL_LINE, IF(name="Quadratic", MODEL_QUAD, IF(name="Cubic", MODEL_CUBIC, IF(name="Exponential", MODEL_EXP, IF(name="Damped Sine", MODEL_DSINE, NA()))))) ) | Get fit model function from its name in the dropdown menu
| MODELJACOBIAN | LAMBDA(name, IF(name="Line", MODEL_LINE_JAC, IF(name="Quadratic", MODEL_QUAD_JAC, IF(name="Cubic", MODEL_CUBIC_JAC, IF(name="Exponential", MODEL_EXP_JAC, IF(name="Damped Sine", MODEL_DSINE_JAC, NA()))))) ) | Get jacobian function from the model name
Third Party Functions - Developed by Vertex42 (MIT License)
| ONES | LAMBDA(m_rows,[n_columns], LET(doc,"https://www.vertex42.com/lambda/ones-and-zeros.html", rows,ROWS(m_rows),cols,COLUMNS(m_rows), IF(OR(rows>1,cols>1), SEQUENCE(rows,cols,1,0), IF(ISOMITTED(n_columns), SEQUENCE(m_rows,1,1,0), SEQUENCE(m_rows,n_columns,1,0) ) ) )) | Generate matrix of ones.
| DIAG | LAMBDA(array,[k], LET(doc,"https://www.vertex42.com/lambda/diag.html", version,"1/2/2026 - Added [k] option", nR, ROWS(array), nC, COLUMNS(array), k, IF(ISOMITTED(k),0,k), isVector, OR(nR=1, nC=1), IF(isVector, LET( v, TOCOL(array), nV, ROWS(v), m, nV+ABS(k), MAKEARRAY(m,m,LAMBDA(i,j, IF( j-i = k, LET(p, IF(k>=0,i,j), IF(AND(p>=1,p<=nV),INDEX(v,p),0) ), 0 ) )) ), LET( nK, IF(k>=0, MAX(0, MIN(nR, nC-k)), MAX(0, MIN(nR+k, nC))), i_0, IF(k>=0,1,1-k), j_0, IF(k>=0,1+k,1), IF(nK=0, NA(), MAKEARRAY(nK,1,LAMBDA(t,_,INDEX(array,i_0+t-1,j_0+t-1))) ) ) ) )) | Create diagonal matrix from array.
| LINSPACE | LAMBDA(start,end,n, LET(doc,"https://www.vertex42.com/lambda/linspace.html", SEQUENCE(n,1,start,(end-start)/(n-1)) )) | Generate sequence of evenly spaced values between start and end points.