## Instructions
Sandia Primary Standards Lab - Smart Measurement Assurance Plan - Risk Calculator
v0.1
uncertainty@sandia.gov
This workbook calculates average and specific measurement decision risk. No Excel Macros are used, however the calculations make use of Excel's LAMBDA function which is only available in Office 365 or 2024.
Spreadsheet Tabs
Risk Calculator - Normal | Calculate average (global) risk assuming a normal distribution of products
Risk Calculator - Gamma | Calculate average (global) risk assuming a gamma distribution of products
Specific Risk | Plot probability of conformance (specific risk) over a range of measured values
Guardbands | Plot probability of false accept and reject (PFA and PFR) as a function of guardband width for a measurement
Risk Curves | Plot PFA versus in-tolerance probability for various TUR, with optional guardbanding
Cost Model | Compare financial tradeoff between PFA and PFR
Lookup | Provides allowable options for dropdown menus. Hidden by default.
About
Sandia National Laboratories is a multimission laboratory managed and operated by National Technology & Engineering Solutions of Sandia, LLC, a wholly owned subsidiary of Honeywell International Inc., for the U.S. Department of Energy’s National Nuclear Security Administration under contract DE-NA0003525.
SAND2025-10736O

## Risk Calculator - Normal
Global (Average) Risk Calculator - Normal Product Distribution | [8 empty] | ← Intermediate and plot data in hidden columns here
Tolerance | 200 | ± | 5 | Lower Tolerance | 195
Nominal/Expected Value | 200 |  |  | Upper Tolerance | 205
In-Tolerance Probability | 85% |  |  | Process Stdev | 3.4734
Measurement Uncertainty (k=2) ± | 2 |  |  | Guardband Width | 0.66127
Guardband Method | PFA |  |  | Lower Acceptance Limit | 195.6613
Manual Guardband Width | 0.1 |  |  | Upper Acceptance Limit | 204.3387
PFA Target | 1%
Acceptance Limit | 200 | ± | 4.3387
TUR | 2.5
Average PFA (No Guardband) | 2.516%
Average PFA(Guardbanded) | 1.011%
Average PFR (Guardbanded) | 9.021%
Worst-case Specific Risk | 25.42%
Process Yield | 85%
Process Risk | 15%
Process Capability, Cpk | 0.47984
|  |  |  | 190 | TL | TU | AL | AU
|  |  |  | 210 | 195 | 205 | 195.6613 | 204.3387
[5 empty] | 195 | 205 | 195.6613 | 204.3387
N
1000
[5 empty] | Product Sample | Measurement Sample
[5 empty] | 198.1506 | 199.0734
[5 empty] | 196.9918 | 197.4236
[5 empty] | 199.5311 | 198.2507
[5 empty] | 197.6762 | 197.5085
[5 empty] | 199.7122 | 200.1282
[5 empty] | 206.0588 | 206.1951
[5 empty] | 193.8798 | 193.0272
[5 empty] | 196.6224 | 197.5856
[5 empty] | 202.0673 | 203.2743
[5 empty] | 196.2312 | 195.0646
[5 empty] | 198.6252 | 199.7862
[5 empty] | 201.4138 | 200.5629
[5 empty] | 200.1362 | 199.7414
[5 empty] | 199.9215 | 201.5125
[5 empty] | 203.1284 | 204.8038
[5 empty] | 201.6432 | 201.9967
[5 empty] | 193.3754 | 194.1696
[5 empty] | 204.9486 | 203.0284
[5 empty] | 195.1397 | 195.2906
[5 empty] | 201.885 | 201.2577
[5 empty] | 202.9785 | 203.508
[5 empty] | 205.9655 | 204.8475
[5 empty] | 200.9256 | 199.3445
[5 empty] | 198.6128 | 197.3862
[5 empty] | 201.8261 | 202.3582
[5 empty] | 204.394 | 202.7439
[5 empty] | 199.8881 | 201.2568
[5 empty] | 199.5369 | 198.0042
[5 empty] | 202.0425 | 202.2904
[5 empty] | 202.5347 | 201.7326
[5 empty] | 202.4741 | 202.4881
[5 empty] | 198.9187 | 198.0481
[5 empty] | 207.0297 | 207.5186
[5 empty] | 199.7171 | 201.2435
[5 empty] | 196.5352 | 194.6521
[5 empty] | 201.2189 | 200.9293
[5 empty] | 195.531 | 195.6055
[5 empty] | 200.9889 | 199.8356
[5 empty] | 197.2285 | 198.0405
[5 empty] | 196.3649 | 195.7288
[5 empty] | 198.3841 | 197.6687
[5 empty] | 197.1879 | 197.0205
[5 empty] | 199.433 | 199.5624
[5 empty] | 195.6147 | 197.7058
[5 empty] | 202.904 | 202.0324
[5 empty] | 198.9349 | 200.4592
[5 empty] | 197.8972 | 197.145
[5 empty] | 203.0331 | 202.8853
[5 empty] | 197.2851 | 198.194
[5 empty] | 206.0029 | 205.8072
[5 empty] | 200.4789 | 201.0808
[5 empty] | 198.7929 | 199.0881
[5 empty] | 198.167 | 198.4426
[5 empty] | 200.2808 | 202.0962
[5 empty] | 199.6372 | 198.7424
[5 empty] | 197.21 | 196.9907
[5 empty] | 198.5584 | 198.737
[5 empty] | 198.1428 | 198.4837
[5 empty] | 203.6211 | 202.9833
[5 empty] | 193.6064 | 193.6872
[5 empty] | 200.9862 | 203.1595
[5 empty] | 195.7515 | 196.4403
[5 empty] | 198.8661 | 198.1371
[5 empty] | 195.0616 | 196.371
[5 empty] | 197.1606 | 197.4554
[5 empty] | 200.7735 | 200.2156
[5 empty] | 200.3082 | 200.2016
[5 empty] | 200.8966 | 201.9027
[5 empty] | 198.9255 | 201.0539
[5 empty] | 199.268 | 199.3555
[5 empty] | 199.6618 | 200.8331
[5 empty] | 197.3027 | 197.8287
[5 empty] | 198.6388 | 200.1365
[5 empty] | 196.1637 | 194.9924
[5 empty] | 202.5809 | 201.8335
[5 empty] | 199.2622 | 198.7246
[5 empty] | 199.3813 | 201.1663
[5 empty] | 198.859 | 197.9367
[5 empty] | 199.8634 | 199.9072
[5 empty] | 202.2041 | 201.0097
[5 empty] | 200.4622 | 201.0989
[5 empty] | 203.6941 | 204.0531
[5 empty] | 198.5239 | 196.4185
[5 empty] | 201.4148 | 202.0581
[5 empty] | 197.422 | 197.732
[5 empty] | 197.5932 | 196.0188
[5 empty] | 198.4572 | 197.5105
[5 empty] | 201.4668 | 202.2495
[5 empty] | 200.9391 | 200.7034
[5 empty] | 197.9776 | 199.4155
[5 empty] | 205.3091 | 206.276
[5 empty] | 202.5843 | 200.9141
[5 empty] | 195.6196 | 195.6123
[5 empty] | 195.4152 | 195.6893
[5 empty] | 194.5498 | 193.7337
[5 empty] | 202.2522 | 202.2448
[5 empty] | 198.9722 | 199.6007
[5 empty] | 198.5978 | 197.8736
[5 empty] | 198.0024 | 199.0135
[5 empty] | 205.0337 | 205.9467
[5 empty] | 199.168 | 198.6945
[5 empty] | 193.6149 | 192.7753
[5 empty] | 199.3856 | 199.7811
[5 empty] | 194.8671 | 193.7461
[5 empty] | 200.3241 | 200.407
[5 empty] | 197.5299 | 198.509
[5 empty] | 192.7193 | 193.8762
[5 empty] | 197.0348 | 195.7575
[5 empty] | 200.0875 | 198.9994
[5 empty] | 205.3608 | 207.6304
[5 empty] | 198.2339 | 198.108
[5 empty] | 198.463 | 198.058
[5 empty] | 199.7499 | 198.5243
[5 empty] | 197.3094 | 197.4538
[5 empty] | 200.5932 | 200.4772
[5 empty] | 199.355 | 200.5231
[5 empty] | 194.3481 | 193.7693
[5 empty] | 192.8678 | 195.3973
[5 empty] | 203.85 | 202.3958
[5 empty] | 199.1218 | 200.3365
[5 empty] | 201.8385 | 202.572
[5 empty] | 206.9609 | 206.1609
[5 empty] | 201.4389 | 201.6695
[5 empty] | 204.0365 | 204.4314
[5 empty] | 196.2586 | 196.7723
[5 empty] | 198.7587 | 196.5869
[5 empty] | 203.398 | 203.8378
[5 empty] | 200.3407 | 198.9932
[5 empty] | 199.1075 | 199.3666
[5 empty] | 198.2508 | 198.2467
[5 empty] | 197.3471 | 197.588
[5 empty] | 200.7219 | 200.5657
[5 empty] | 199.9855 | 200.7221
[5 empty] | 202.923 | 204.0364
[5 empty] | 197.9208 | 197.3243
[5 empty] | 203.1392 | 202.9385
[5 empty] | 196.5176 | 195.9215
[5 empty] | 199.7941 | 198.895
[5 empty] | 202.8725 | 202.3445
[5 empty] | 198.0574 | 197.7757
[5 empty] | 202.3992 | 202.3086
[5 empty] | 197.3024 | 196.8408
[5 empty] | 204.1069 | 204.4286
[5 empty] | 195.4712 | 195.2256
[5 empty] | 198.0591 | 198.4932
[5 empty] | 199.8932 | 200.2972
[5 empty] | 204.9795 | 205.7019
[5 empty] | 197.5387 | 197.7957
[5 empty] | 194.5242 | 191.9689
[5 empty] | 202.3023 | 203.1467
[5 empty] | 194.8275 | 194.2656
[5 empty] | 199.7165 | 199.5164
[5 empty] | 201.8496 | 201.0869
[5 empty] | 201.5886 | 200.41
[5 empty] | 202.5709 | 201.2238
[5 empty] | 201.8393 | 201.9778
[5 empty] | 192.6641 | 193.8933
[5 empty] | 200.6434 | 199.4135
[5 empty] | 202.2775 | 202.6569
[5 empty] | 203.7728 | 204.4694
[5 empty] | 199.7508 | 200.4875
[5 empty] | 201.4066 | 200.8524
[5 empty] | 207.2579 | 207.9433
[5 empty] | 199.3763 | 196.9212
[5 empty] | 201.7814 | 201.727
[5 empty] | 202.4973 | 201.6628
[5 empty] | 205.1972 | 205.174
[5 empty] | 198.0601 | 196.9819
[5 empty] | 203.6163 | 204.3179
[5 empty] | 195.4761 | 194.9175
[5 empty] | 206.7976 | 206.4079
[5 empty] | 200.7667 | 201.6976
[5 empty] | 200.2364 | 201.4669
[5 empty] | 196.6678 | 197.3295
[5 empty] | 201.6661 | 204.1196
[5 empty] | 194.3404 | 194.49
[5 empty] | 203.9681 | 202.4742
[5 empty] | 201.0675 | 203.1625
[5 empty] | 204.1627 | 204.1451
[5 empty] | 194.5462 | 195.2695
[5 empty] | 202.5042 | 201.2167
[5 empty] | 197.7262 | 197.4169
[5 empty] | 204.7885 | 203.3789
[5 empty] | 196.9824 | 197.7923
[5 empty] | 198.3173 | 198.3747
[5 empty] | 200.8837 | 202.8526
[5 empty] | 200.5342 | 202.365
[5 empty] | 199.738 | 200.516
[5 empty] | 201.0002 | 201.0008
[5 empty] | 200.19 | 200.2124
[5 empty] | 196.3039 | 196.8578
[5 empty] | 205.7674 | 205.7126
[5 empty] | 197.4512 | 197.1416
[5 empty] | 198.6736 | 200.0139
[5 empty] | 196.9857 | 195.7891
[5 empty] | 193.309 | 191.3755
[5 empty] | 196.4636 | 197.21
[5 empty] | 195.3779 | 197.0498
[5 empty] | 201.3167 | 201.6144
[5 empty] | 199.5166 | 198.2613
[5 empty] | 199.6153 | 200.9854
[5 empty] | 196.9705 | 194.8308
[5 empty] | 198.9303 | 197.7406
[5 empty] | 203.3666 | 203.2816
[5 empty] | 204.3757 | 204.9322
[5 empty] | 195.5279 | 194.1808
[5 empty] | 204.1575 | 201.8854
[5 empty] | 194.0381 | 194.8233
[5 empty] | 200.7597 | 201.6356
[5 empty] | 198.0926 | 197.8705
[5 empty] | 193.157 | 190.603
[5 empty] | 202.8243 | 203.4527
[5 empty] | 198.619 | 197.9846
[5 empty] | 198.9597 | 199.6816
[5 empty] | 199.525 | 198.007
[5 empty] | 199.0537 | 200.0591
[5 empty] | 200.6987 | 202.0893
[5 empty] | 198.383 | 198.5888
[5 empty] | 199.7634 | 199.0534
[5 empty] | 200.536 | 200.2367
[5 empty] | 202.7595 | 202.5835
[5 empty] | 200.3494 | 200.1288
[5 empty] | 198.5563 | 199.064
[5 empty] | 201.2578 | 203.6975
[5 empty] | 199.2156 | 199.5439
[5 empty] | 202.4713 | 203.1132
[5 empty] | 197.1877 | 196.018
[5 empty] | 201.9816 | 201.3545
[5 empty] | 204.6385 | 205.4019
[5 empty] | 203.2578 | 202.2822
[5 empty] | 198.2123 | 198.4118
[5 empty] | 196.4397 | 195.8257
[5 empty] | 201.1507 | 201.0711
[5 empty] | 203.6514 | 203.6391
[5 empty] | 198.7283 | 197.7039
[5 empty] | 198.7983 | 197.5071
[5 empty] | 201.484 | 199.8164
[5 empty] | 195.4889 | 194.9379
[5 empty] | 197.8598 | 196.4739
[5 empty] | 198.8582 | 196.7406
[5 empty] | 198.4871 | 198.8016
[5 empty] | 203.4678 | 201.9769
[5 empty] | 199.2888 | 199.0719
[5 empty] | 197.8716 | 196.6261
[5 empty] | 193.948 | 193.6153
[5 empty] | 198.4565 | 197.0079
[5 empty] | 199.4975 | 200.2899
[5 empty] | 197.8204 | 198.5289
[5 empty] | 207.7955 | 208.7526
[5 empty] | 198.9574 | 199.329
[5 empty] | 200.1704 | 200.2598
[5 empty] | 205.0745 | 203.5229
[5 empty] | 198.698 | 198.2775
[5 empty] | 197.805 | 198.1933
[5 empty] | 201.8843 | 202.8116
[5 empty] | 205.1165 | 205.7783
[5 empty] | 201.757 | 201.9525
[5 empty] | 193.303 | 192.6931
[5 empty] | 202.7227 | 201.5036
[5 empty] | 205.7749 | 206.0346
[5 empty] | 199.1748 | 198.5704
[5 empty] | 201.743 | 202.9708
[5 empty] | 198.2303 | 198.1332
[5 empty] | 203.2549 | 203.4483
[5 empty] | 198.8694 | 199.7475
[5 empty] | 200.2091 | 199.4686
[5 empty] | 193.1737 | 193.8509
[5 empty] | 201.7318 | 200.4987
[5 empty] | 199.445 | 199.8776
[5 empty] | 194.1152 | 195.1139
[5 empty] | 201.3129 | 200.7972
[5 empty] | 191.3321 | 190.3696
[5 empty] | 203.7025 | 202.6575
[5 empty] | 200.058 | 201.605
[5 empty] | 198.7465 | 198.5243
[5 empty] | 200.9556 | 200.4963
[5 empty] | 197.2684 | 196.3587
[5 empty] | 200.8034 | 199.5557
[5 empty] | 200.8334 | 199.8244
[5 empty] | 198.0015 | 197.6297
[5 empty] | 196.0055 | 195.8785
[5 empty] | 200.66 | 199.8437
[5 empty] | 203.3658 | 201.9769
[5 empty] | 196.2824 | 196.7594
[5 empty] | 197.1345 | 197.4431
[5 empty] | 200.145 | 201.1472
[5 empty] | 200.6074 | 202.7121
[5 empty] | 201.4088 | 200.6465
[5 empty] | 200.5122 | 200.4027
[5 empty] | 199.9334 | 199.1735
[5 empty] | 205.8819 | 205.0713
[5 empty] | 193.5811 | 193.7028
[5 empty] | 200.762 | 200.0755
[5 empty] | 200.3843 | 201.0277
[5 empty] | 199.0777 | 199.1655
[5 empty] | 200.9109 | 200.108
[5 empty] | 200.4931 | 199.9154
[5 empty] | 202.2929 | 201.0169
[5 empty] | 202.6891 | 204.1661
[5 empty] | 201.791 | 200.3045
[5 empty] | 200.2804 | 199.8271
[5 empty] | 200.2979 | 202.2291
[5 empty] | 197.3677 | 197.5243
[5 empty] | 197.5374 | 199.2912
[5 empty] | 201.9652 | 201.6161
[5 empty] | 201.0847 | 201.832
[5 empty] | 196.7349 | 195.8197
[5 empty] | 197.1518 | 197.0978
[5 empty] | 205.2048 | 204.5093
[5 empty] | 203.0349 | 203.8235
[5 empty] | 194.2139 | 192.852
[5 empty] | 196.3611 | 196.0818
[5 empty] | 204.2444 | 203.1117
[5 empty] | 199.0781 | 200.6098
[5 empty] | 196.9247 | 196.5097
[5 empty] | 201.8135 | 201.7566
[5 empty] | 200.0195 | 199.7977
[5 empty] | 203.2727 | 201.8349
[5 empty] | 205.7971 | 206.0267
[5 empty] | 200.1809 | 200.8103
[5 empty] | 191.9286 | 189.5445
[5 empty] | 205.6647 | 205.332
[5 empty] | 202.6811 | 202.6874
[5 empty] | 202.3505 | 202.5919
[5 empty] | 198.1145 | 200.1673
[5 empty] | 202.7183 | 204.2799
[5 empty] | 199.1289 | 197.8081
[5 empty] | 197.8942 | 199.3447
[5 empty] | 197.8246 | 195.6278
[5 empty] | 207.512 | 206.8738
[5 empty] | 198.9139 | 197.1338
[5 empty] | 202.2112 | 200.8884
[5 empty] | 195.8564 | 195.6363
[5 empty] | 200.3175 | 198.7179
[5 empty] | 207.4604 | 208.4567
[5 empty] | 195.7596 | 196.3433
[5 empty] | 199.5764 | 198.8152
[5 empty] | 202.6255 | 203.3032
[5 empty] | 196.2914 | 197.8472
[5 empty] | 198.4392 | 198.9561
[5 empty] | 198.559 | 199.7098
[5 empty] | 196.4212 | 197.4708
[5 empty] | 202.2285 | 205.258
[5 empty] | 199.6571 | 199.7515
[5 empty] | 194.1424 | 194.791
[5 empty] | 204.5608 | 205.18
[5 empty] | 201.1937 | 203.1603
[5 empty] | 195.1205 | 195.985
[5 empty] | 205.4969 | 205.6209
[5 empty] | 199.3272 | 200.8565
[5 empty] | 200.5589 | 198.8738
[5 empty] | 188.8724 | 189.681
[5 empty] | 202.4279 | 203.9032
[5 empty] | 202.5566 | 200.2639
[5 empty] | 202.6892 | 202.6227
[5 empty] | 199.2003 | 198.8604
[5 empty] | 195.731 | 194.1165
[5 empty] | 198.036 | 198.0814
[5 empty] | 201.9286 | 201.7894
[5 empty] | 204.3649 | 204.9311
[5 empty] | 193.2922 | 194.3572
[5 empty] | 206.2171 | 207.0762
[5 empty] | 196.5757 | 197.6923
[5 empty] | 201.4719 | 202.5804
[5 empty] | 204.252 | 204.2199
[5 empty] | 200.8223 | 200.3822
[5 empty] | 199.7379 | 199.8285
[5 empty] | 198.4845 | 199.4101
[5 empty] | 200.5591 | 199.9871
[5 empty] | 200.1618 | 201.1175
[5 empty] | 199.6819 | 199.601
[5 empty] | 203.3824 | 202.0548
[5 empty] | 205.5773 | 204.7696
[5 empty] | 200.3417 | 200.0815
[5 empty] | 200.6185 | 198.8127
[5 empty] | 200.3044 | 199.627
[5 empty] | 199.0728 | 199.5021
[5 empty] | 195.6896 | 194.9492
[5 empty] | 195.9315 | 194.7613
[5 empty] | 201.0047 | 203.189
[5 empty] | 202.7995 | 200.5848
[5 empty] | 201.7066 | 201.6201
[5 empty] | 206.4165 | 207.6292
[5 empty] | 201.1583 | 198.9952
[5 empty] | 194.1572 | 194.7996
[5 empty] | 202.7099 | 203.4114
[5 empty] | 202.2261 | 202.6317
[5 empty] | 204.6477 | 203.858
[5 empty] | 203.2585 | 204.0435
[5 empty] | 200.4762 | 200.2649
[5 empty] | 197.8343 | 196.5809
[5 empty] | 199.2175 | 200.9202
[5 empty] | 201.1853 | 201.4366
[5 empty] | 202.8675 | 202.4741
[5 empty] | 200.8353 | 202.0395
[5 empty] | 205.208 | 206.1818
[5 empty] | 201.5237 | 201.3932
[5 empty] | 196.1041 | 196.149
[5 empty] | 199.0114 | 198.4405
[5 empty] | 204.9413 | 204.6263
[5 empty] | 202.1434 | 201.1554
[5 empty] | 201.8011 | 202.1092
[5 empty] | 204.6768 | 204.0968
[5 empty] | 204.74 | 204.4767
[5 empty] | 201.1687 | 201.3992
[5 empty] | 196.7624 | 197.9437
[5 empty] | 196.1966 | 196.5278
[5 empty] | 200.5258 | 198.4055
[5 empty] | 202.8638 | 203.5487
[5 empty] | 198.3925 | 196.6944
[5 empty] | 193.7441 | 193.2915
[5 empty] | 203.6278 | 204.4713
[5 empty] | 195.3041 | 195.9409
[5 empty] | 196.5567 | 196.9583
[5 empty] | 200.3912 | 199.4387
[5 empty] | 198.2302 | 198.2255
[5 empty] | 203.9879 | 204.746
[5 empty] | 200.7363 | 201.039
[5 empty] | 200.4184 | 202.1547
[5 empty] | 197.4367 | 197.161
[5 empty] | 197.2456 | 197.6198
[5 empty] | 202.4438 | 202.0451
[5 empty] | 201.5652 | 199.8502
[5 empty] | 202.2193 | 202.2104
[5 empty] | 202.6175 | 203.6177
[5 empty] | 201.2091 | 200.8294
[5 empty] | 201.2098 | 200.371
[5 empty] | 204.6886 | 203.3481
[5 empty] | 196.9736 | 197.0818
[5 empty] | 204.3253 | 204.0497
[5 empty] | 197.2901 | 198.0359
[5 empty] | 200.8096 | 200.8167
[5 empty] | 196.2774 | 196.9166
[5 empty] | 198.0629 | 198.5893
[5 empty] | 197.9394 | 199.5145
[5 empty] | 202.061 | 202.9501
[5 empty] | 200.4388 | 200.8124
[5 empty] | 194.2594 | 195.3358
[5 empty] | 200.7033 | 200.3219
[5 empty] | 192.815 | 190.9999
[5 empty] | 199.6953 | 200.2835
[5 empty] | 197.6269 | 197.8382
[5 empty] | 204.7289 | 203.1121
[5 empty] | 205.6404 | 205.8243
[5 empty] | 204.4437 | 205.2424
[5 empty] | 206.2117 | 206.0073
[5 empty] | 203.1662 | 202.9455
[5 empty] | 203.7161 | 205.1382
[5 empty] | 206.2686 | 207.056
[5 empty] | 203.288 | 202.8869
[5 empty] | 200.7777 | 201.3498
[5 empty] | 200.9458 | 201.5733
[5 empty] | 198.6078 | 199.0436
[5 empty] | 197.1001 | 196.8682
[5 empty] | 202.8019 | 203.6194
[5 empty] | 202.9591 | 201.2607
[5 empty] | 203.3327 | 203.2423
[5 empty] | 200.2381 | 199.4041
[5 empty] | 192.9749 | 193.0977
[5 empty] | 200.3016 | 201.2124
[5 empty] | 195.9576 | 196.4019
[5 empty] | 197.5593 | 197.2051
[5 empty] | 198.6438 | 199.635
[5 empty] | 200.2189 | 199.1352
[5 empty] | 193.6907 | 193.4783
[5 empty] | 202.0558 | 201.2125
[5 empty] | 200.4827 | 200.3939
[5 empty] | 199.5889 | 199.7125
[5 empty] | 197.7361 | 198.5067
[5 empty] | 203.5338 | 203.5164
[5 empty] | 203.0856 | 204.6647
[5 empty] | 201.3713 | 201.9244
[5 empty] | 210.0657 | 211.0019
[5 empty] | 201.527 | 201.3979
[5 empty] | 198.6387 | 199.338
[5 empty] | 200.7656 | 200.8214
[5 empty] | 200.1089 | 200.1846
[5 empty] | 199.5272 | 199.2111
[5 empty] | 201.4043 | 200.8338
[5 empty] | 197.8083 | 197.4669
[5 empty] | 199.7795 | 200.2925
[5 empty] | 195.6502 | 195.197
[5 empty] | 201.5896 | 201.9194
[5 empty] | 206.098 | 206.1282
[5 empty] | 207.4746 | 207.9359
[5 empty] | 200.1011 | 199.7826
[5 empty] | 198.7474 | 197.7186
[5 empty] | 195.314 | 194.5784
[5 empty] | 204.1161 | 203.9447
[5 empty] | 204.0125 | 202.4683
[5 empty] | 201.2396 | 201.5446
[5 empty] | 200.3776 | 201.2248
[5 empty] | 199.0355 | 199.5008
[5 empty] | 196.5023 | 195.6883
[5 empty] | 201.7769 | 202.4157
[5 empty] | 203.1781 | 202.1023
[5 empty] | 203.8606 | 203.044
[5 empty] | 195.8722 | 196.5173
[5 empty] | 197.3323 | 197.7678
[5 empty] | 198.9786 | 197.7021
[5 empty] | 197.8052 | 197.912
[5 empty] | 192.9161 | 192.0604
[5 empty] | 190.1085 | 191.3226
[5 empty] | 201.0525 | 199.6271
[5 empty] | 195.4522 | 195.2163
[5 empty] | 197.6275 | 196.8326
[5 empty] | 199.5047 | 200.7243
[5 empty] | 197.8532 | 196.6034
[5 empty] | 211.5033 | 212.3777
[5 empty] | 195.8969 | 195.7372
[5 empty] | 200.5391 | 201.5751
[5 empty] | 202.219 | 203.4152
[5 empty] | 200.2783 | 200.6082
[5 empty] | 203.2771 | 202.8343
[5 empty] | 200.5277 | 200.8101
[5 empty] | 201.3468 | 202.4022
[5 empty] | 193.7176 | 194.6847
[5 empty] | 199.44 | 201.6246
[5 empty] | 199.35 | 199.5139
[5 empty] | 194.4608 | 193.0542
[5 empty] | 199.022 | 196.8071
[5 empty] | 203.5842 | 204.6697
[5 empty] | 201.5916 | 203.1163
[5 empty] | 200.169 | 199.3305
[5 empty] | 197.1751 | 197.3277
[5 empty] | 199.4508 | 200.8045
[5 empty] | 196.4446 | 195.3654
[5 empty] | 199.5535 | 199.8678
[5 empty] | 203.331 | 202.289
[5 empty] | 202.8642 | 203.0363
[5 empty] | 201.9997 | 201.9942
[5 empty] | 197.797 | 197.8144
[5 empty] | 200.7259 | 200.3818
[5 empty] | 196.7009 | 198.7831
[5 empty] | 200.279 | 200.43
[5 empty] | 192.756 | 192.1977
[5 empty] | 192.872 | 192.0713
[5 empty] | 199.0847 | 198.3004
[5 empty] | 202.2366 | 203.5038
[5 empty] | 199.2343 | 198.9495
[5 empty] | 207.352 | 206.608
[5 empty] | 196.8959 | 198.1245
[5 empty] | 195.8967 | 197.7328
[5 empty] | 201.3539 | 201.0032
[5 empty] | 199.382 | 197.1526
[5 empty] | 202.3751 | 203.0948
[5 empty] | 196.1504 | 198.0743
[5 empty] | 199.8754 | 199.9089
[5 empty] | 197.8032 | 200.0041
[5 empty] | 200.3259 | 200.0335
[5 empty] | 199.4342 | 198.4383
[5 empty] | 202.1733 | 201.6958
[5 empty] | 197.7776 | 198.0946
[5 empty] | 199.6462 | 199.6795
[5 empty] | 202.1565 | 201.732
[5 empty] | 205.3433 | 204.2756
[5 empty] | 195.9393 | 197.4044
[5 empty] | 201.228 | 201.4372
[5 empty] | 198.4726 | 197.5577
[5 empty] | 199.0529 | 198.9381
[5 empty] | 196.6944 | 197.4348
[5 empty] | 194.0973 | 192.9447
[5 empty] | 197.7081 | 198.1232
[5 empty] | 201.4865 | 201.4129
[5 empty] | 197.9466 | 197.1714
[5 empty] | 199.5638 | 200.8017
[5 empty] | 192.858 | 193.0527
[5 empty] | 196.1381 | 196.3186
[5 empty] | 200.4441 | 200.6565
[5 empty] | 195.438 | 196.4699
[5 empty] | 195.5391 | 196.0603
[5 empty] | 202.6341 | 202.0513
[5 empty] | 205.3438 | 206.203
[5 empty] | 198.3214 | 198.722
[5 empty] | 197.1942 | 198.1186
[5 empty] | 204.7286 | 204.0916
[5 empty] | 198.9191 | 200.1156
[5 empty] | 193.6999 | 192.9775
[5 empty] | 197.2401 | 197.3046
[5 empty] | 197.7722 | 198.9588
[5 empty] | 201.0154 | 200.6942
[5 empty] | 201.091 | 201.6606
[5 empty] | 199.8839 | 200.4469
[5 empty] | 201.8242 | 201.4124
[5 empty] | 201.9175 | 201.46
[5 empty] | 199.3378 | 199.0862
[5 empty] | 197.0866 | 195.3984
[5 empty] | 205.1763 | 206.188
[5 empty] | 195.483 | 195.1819
[5 empty] | 199.1848 | 199.3633
[5 empty] | 197.5903 | 195.3991
[5 empty] | 204.7281 | 206.2694
[5 empty] | 198.2615 | 199.1377
[5 empty] | 196.1214 | 194.6203
[5 empty] | 198.2973 | 197.5817
[5 empty] | 200.9935 | 202.0385
[5 empty] | 193.6548 | 192.2394
[5 empty] | 193.9759 | 193.4727
[5 empty] | 196.7265 | 195.5862
[5 empty] | 195.2533 | 196.6349
[5 empty] | 198.2303 | 199.2165
[5 empty] | 202.0938 | 201.9599
[5 empty] | 199.3379 | 198.6712
[5 empty] | 201.2514 | 201.2666
[5 empty] | 201.2956 | 201.9913
[5 empty] | 201.9715 | 204.5393
[5 empty] | 202.3368 | 200.3056
[5 empty] | 198.1097 | 198.5488
[5 empty] | 202.64 | 204.046
[5 empty] | 204.6832 | 204.9296
[5 empty] | 199.6234 | 199.8945
[5 empty] | 195.5799 | 195.632
[5 empty] | 198.4491 | 197.6076
[5 empty] | 202.3633 | 202.3928
[5 empty] | 203.322 | 204.352
[5 empty] | 191.7175 | 191.457
[5 empty] | 192.8657 | 192.1156
[5 empty] | 201.4851 | 200.6791
[5 empty] | 205.7055 | 205.3142
[5 empty] | 198.3742 | 199.9336
[5 empty] | 204.5102 | 203.4934
[5 empty] | 200.6849 | 201.6
[5 empty] | 199.4167 | 199.1259
[5 empty] | 204.7498 | 204.0513
[5 empty] | 202.6738 | 202.1432
[5 empty] | 197.7936 | 198.7289
[5 empty] | 203.5479 | 203.183
[5 empty] | 203.4958 | 202.0392
[5 empty] | 202.0379 | 203.7111
[5 empty] | 205.2388 | 205.2609
[5 empty] | 192.7734 | 192.2831
[5 empty] | 198.0936 | 198.1535
[5 empty] | 198.8913 | 199.4591
[5 empty] | 205.3318 | 205.4883
[5 empty] | 204.2009 | 203.9656
[5 empty] | 190.1548 | 190.1055
[5 empty] | 198.6874 | 197.9904
[5 empty] | 199.0074 | 198.707
[5 empty] | 200.1354 | 200.9367
[5 empty] | 198.5966 | 197.9144
[5 empty] | 199.3579 | 201.5253
[5 empty] | 200.9974 | 200.2433
[5 empty] | 195.2797 | 195.0221
[5 empty] | 202.3175 | 201.4538
[5 empty] | 204.7032 | 205.1761
[5 empty] | 197.1301 | 195.9047
[5 empty] | 199.2893 | 199.4555
[5 empty] | 201.0882 | 201.2842
[5 empty] | 200.7147 | 199.6427
[5 empty] | 198.7345 | 198.5514
[5 empty] | 201.5266 | 201.5262
[5 empty] | 198.1367 | 199.1986
[5 empty] | 197.6164 | 196.3419
[5 empty] | 205.4709 | 205.6196
[5 empty] | 195.0089 | 194.0044
[5 empty] | 195.0767 | 196.1185
[5 empty] | 204.3703 | 202.8496
[5 empty] | 203.0085 | 203.1701
[5 empty] | 197.215 | 196.7101
[5 empty] | 207.3415 | 205.7804
[5 empty] | 200.7342 | 200.4907
[5 empty] | 205.304 | 207.531
[5 empty] | 197.4046 | 196.902
[5 empty] | 206.0498 | 205.4595
[5 empty] | 199.281 | 200.5134
[5 empty] | 196.7772 | 197.6064
[5 empty] | 200.5029 | 200.9417
[5 empty] | 197.1876 | 195.9034
[5 empty] | 197.8438 | 198.4637
[5 empty] | 199.3637 | 198.4253
[5 empty] | 198.7305 | 199.2353
[5 empty] | 198.2024 | 195.8076
[5 empty] | 199.0569 | 200.405
[5 empty] | 203.7382 | 203.8957
[5 empty] | 202.2899 | 201.4479
[5 empty] | 198.6667 | 198.5694
[5 empty] | 208.8547 | 209.6332
[5 empty] | 195.4459 | 195.7473
[5 empty] | 201.8885 | 202.9847
[5 empty] | 199.8987 | 198.7508
[5 empty] | 199.0478 | 201.4246
[5 empty] | 199.6752 | 200.2749
[5 empty] | 195.3237 | 195.2688
[5 empty] | 202.2809 | 201.4362
[5 empty] | 197.5123 | 195.9456
[5 empty] | 201.451 | 202.3727
[5 empty] | 197.39 | 198.515
[5 empty] | 202.1313 | 202.8366
[5 empty] | 196.5692 | 197.5004
[5 empty] | 203.161 | 201.8408
[5 empty] | 207.0928 | 206.52
[5 empty] | 200.6133 | 199.2082
[5 empty] | 203.4226 | 204.0073
[5 empty] | 203.4377 | 202.0879
[5 empty] | 196.889 | 196.594
[5 empty] | 202.0148 | 203.1087
[5 empty] | 197.4103 | 197.6406
[5 empty] | 197.6498 | 197.3999
[5 empty] | 195.8914 | 196.0549
[5 empty] | 205.6162 | 206.3735
[5 empty] | 196.1267 | 196.8768
[5 empty] | 196.8024 | 198.6284
[5 empty] | 196.9855 | 197.7174
[5 empty] | 204.7771 | 205.0736
[5 empty] | 197.8877 | 197.3452
[5 empty] | 199.8545 | 199.3786
[5 empty] | 200.3688 | 199.6734
[5 empty] | 197.8437 | 198.8088
[5 empty] | 191.0524 | 190.6033
[5 empty] | 202.0902 | 200.2876
[5 empty] | 196.4683 | 195.5694
[5 empty] | 203.8239 | 203.9805
[5 empty] | 201.9068 | 201.8254
[5 empty] | 196.9713 | 197.5082
[5 empty] | 197.7051 | 198.865
[5 empty] | 202.7635 | 204.2329
[5 empty] | 202.3333 | 203.2441
[5 empty] | 207.0614 | 207.3072
[5 empty] | 198.2424 | 198.4527
[5 empty] | 195.0403 | 195.9369
[5 empty] | 205.6013 | 205.8576
[5 empty] | 200.7826 | 200.7558
[5 empty] | 190.1171 | 190.5775
[5 empty] | 205.0163 | 205.1902
[5 empty] | 199.469 | 198.8545
[5 empty] | 195.5416 | 195.5413
[5 empty] | 197.5984 | 197.6924
[5 empty] | 206.3118 | 205.7031
[5 empty] | 201.8398 | 201.4805
[5 empty] | 200.7561 | 202.3607
[5 empty] | 201.4722 | 200.8597
[5 empty] | 195.9071 | 194.8264
[5 empty] | 205.9229 | 205.2418
[5 empty] | 203.2104 | 203.84
[5 empty] | 194.2861 | 192.2131
[5 empty] | 199.9658 | 199.7771
[5 empty] | 201.9977 | 202.055
[5 empty] | 206.3718 | 204.3707
[5 empty] | 201.5386 | 200.3282
[5 empty] | 198.8759 | 198.7844
[5 empty] | 193.2374 | 192.0914
[5 empty] | 191.6733 | 193.3259
[5 empty] | 204.059 | 203.8331
[5 empty] | 198.2005 | 198.2489
[5 empty] | 200.137 | 200.8643
[5 empty] | 198.98 | 199.2278
[5 empty] | 194.9523 | 195.3208
[5 empty] | 199.4577 | 199.6941
[5 empty] | 199.1229 | 199.9573
[5 empty] | 197.3748 | 195.9044
[5 empty] | 200.3587 | 199.8621
[5 empty] | 195.8584 | 196.1842
[5 empty] | 201.6271 | 200.5905
[5 empty] | 200.6155 | 200.295
[5 empty] | 201.3544 | 201.4093
[5 empty] | 202.2739 | 203.2596
[5 empty] | 199.2702 | 199.5334
[5 empty] | 199.2458 | 200.1195
[5 empty] | 205.5881 | 205.9596
[5 empty] | 204.9726 | 204.2324
[5 empty] | 204.6441 | 206.1097
[5 empty] | 203.6177 | 203.5616
[5 empty] | 200.1488 | 201.8153
[5 empty] | 202.6385 | 203.0618
[5 empty] | 202.7159 | 204.6862
[5 empty] | 200.6875 | 200.8007
[5 empty] | 196.825 | 196.4867
[5 empty] | 194.0074 | 196.0234
[5 empty] | 198.9717 | 196.8447
[5 empty] | 203.866 | 204.9297
[5 empty] | 197.2992 | 197.3099
[5 empty] | 201.7513 | 200.3783
[5 empty] | 203.9228 | 204.6037
[5 empty] | 201.9128 | 203.4168
[5 empty] | 198.8426 | 198.7549
[5 empty] | 192.7598 | 192.0759
[5 empty] | 201.1518 | 201.1686
[5 empty] | 203.7194 | 203.7131
[5 empty] | 200.1851 | 199.6534
[5 empty] | 197.0556 | 198.7624
[5 empty] | 208.0851 | 208.041
[5 empty] | 194.5118 | 194.1865
[5 empty] | 197.3336 | 196.5046
[5 empty] | 197.2481 | 197.7086
[5 empty] | 199.12 | 199.2725
[5 empty] | 199.9528 | 201.0117
[5 empty] | 203.2514 | 204.4453
[5 empty] | 199.9836 | 199.7685
[5 empty] | 192.7366 | 192.7062
[5 empty] | 197.9666 | 199.2625
[5 empty] | 200.552 | 201.9064
[5 empty] | 195.5538 | 193.9484
[5 empty] | 200.0762 | 198.6418
[5 empty] | 196.7756 | 197.1371
[5 empty] | 191.0544 | 191.768
[5 empty] | 206.2631 | 204.8394
[5 empty] | 197.2285 | 195.5455
[5 empty] | 204.3107 | 204.091
[5 empty] | 193.2974 | 191.3787
[5 empty] | 200.4551 | 201.1714
[5 empty] | 199.6823 | 198.7494
[5 empty] | 201.2128 | 200.816
[5 empty] | 207.0212 | 208.9218
[5 empty] | 195.3502 | 194.3007
[5 empty] | 199.4938 | 199.0916
[5 empty] | 195.1415 | 194.4168
[5 empty] | 197.9767 | 198.568
[5 empty] | 196.5537 | 195.9333
[5 empty] | 202.7389 | 203.4875
[5 empty] | 193.4955 | 193.0537
[5 empty] | 201.5046 | 201.6115
[5 empty] | 196.7207 | 196.8126
[5 empty] | 203.1788 | 203.2597
[5 empty] | 204.5663 | 207.3499
[5 empty] | 200.2843 | 199.2034
[5 empty] | 198.9479 | 201.108
[5 empty] | 204.0458 | 202.6973
[5 empty] | 204.0289 | 204.9115
[5 empty] | 203.9543 | 205.3086
[5 empty] | 202.388 | 201.3645
[5 empty] | 202.0063 | 201.5431
[5 empty] | 201.4823 | 201.8233
[5 empty] | 197.99 | 198.3798
[5 empty] | 196.846 | 198.0606
[5 empty] | 200.9494 | 201.9073
[5 empty] | 201.9498 | 201.9461
[5 empty] | 197.1271 | 196.8908
[5 empty] | 200.8383 | 202.3479
[5 empty] | 201.9132 | 202.5487
[5 empty] | 197.5479 | 198.4294
[5 empty] | 200.1425 | 201.4575
[5 empty] | 195.7048 | 195.0816
[5 empty] | 207.4316 | 207.2165
[5 empty] | 203.9735 | 205.3201
[5 empty] | 198.8036 | 198.8054
[5 empty] | 199.9987 | 199.8362
[5 empty] | 199.0637 | 198.4518
[5 empty] | 207.7251 | 208.0393
[5 empty] | 199.4563 | 198.5392
[5 empty] | 201.4674 | 200.961
[5 empty] | 198.2252 | 196.7707
[5 empty] | 201.1694 | 201.7138
[5 empty] | 203.2136 | 203.1369
[5 empty] | 198.0923 | 198.8278
[5 empty] | 200.2668 | 199.5338
[5 empty] | 207.4623 | 208.0965
[5 empty] | 203.657 | 202.1133
[5 empty] | 200.0778 | 200.1997
[5 empty] | 200.5492 | 200.1134
[5 empty] | 194.6987 | 193.1219
[5 empty] | 196.7689 | 197.1065
[5 empty] | 196.6373 | 196.8302
[5 empty] | 200.183 | 201.5027
[5 empty] | 193.2574 | 192.4824
[5 empty] | 200.7052 | 199.523
[5 empty] | 207.115 | 205.3325
[5 empty] | 202.1341 | 200.502
[5 empty] | 197.563 | 197.5656
[5 empty] | 195.5973 | 196.2772
[5 empty] | 200.1513 | 199.6251
[5 empty] | 199.1215 | 198.2709
[5 empty] | 196.9454 | 196.889
[5 empty] | 200.7744 | 201.1099
[5 empty] | 203.0085 | 202.839
[5 empty] | 197.6601 | 197.3521
[5 empty] | 192.1256 | 193.9725
[5 empty] | 196.6121 | 197.0577
[5 empty] | 195.2424 | 195.7558
[5 empty] | 201.5878 | 199.2198
[5 empty] | 198.9995 | 199.2913
[5 empty] | 199.8352 | 198.9435
[5 empty] | 196.9784 | 196.0749
[5 empty] | 198.4677 | 200.4409
[5 empty] | 198.0644 | 199.6415
[5 empty] | 203.2864 | 203.2749
[5 empty] | 197.2764 | 198.683
[5 empty] | 201.8873 | 202.7668
[5 empty] | 202.1923 | 200.8192
[5 empty] | 200.5314 | 200.753
[5 empty] | 201.2529 | 201.4784
[5 empty] | 202.3486 | 202.5763
[5 empty] | 199.137 | 199.7072
[5 empty] | 203.657 | 202.5267
[5 empty] | 197.2504 | 196.3475
[5 empty] | 199.9357 | 200.8934
[5 empty] | 195.2848 | 196.6945
[5 empty] | 196.6923 | 198.1589
[5 empty] | 204.7885 | 203.7734
[5 empty] | 202.4153 | 201.4392
[5 empty] | 201.4969 | 201.061
[5 empty] | 195.9613 | 195.1495
[5 empty] | 201.7438 | 200.2529
[5 empty] | 199.5674 | 198.4992
[5 empty] | 201.2294 | 202.9374
[5 empty] | 197.0775 | 197.2046
[5 empty] | 200.5373 | 200.5819
[5 empty] | 200.9804 | 200.9331
[5 empty] | 202.8315 | 203.4946
[5 empty] | 201.6875 | 201.6867
[5 empty] | 199.9688 | 200.1533
[5 empty] | 202.8815 | 202.0481
[5 empty] | 200.9372 | 201.2866
[5 empty] | 196.5376 | 196.8133
[5 empty] | 200.714 | 201.5837
[5 empty] | 197.4604 | 198.0933
[5 empty] | 199.8709 | 198.7783
[5 empty] | 199.6739 | 200.9884
[5 empty] | 201.1244 | 202.2634
[5 empty] | 197.5819 | 198.1198
[5 empty] | 195.9396 | 195.404
[5 empty] | 205.5504 | 204.9744
[5 empty] | 206.86 | 206.3723
[5 empty] | 199.8904 | 199.9283
[5 empty] | 198.6237 | 199.2404
[5 empty] | 200.5634 | 199.4076
[5 empty] | 195.5256 | 195.2616
[5 empty] | 197.7664 | 197.1576
[5 empty] | 204.3633 | 204.3567
[5 empty] | 204.9829 | 207.0359
[5 empty] | 200.3602 | 199.0443
[5 empty] | 206.6099 | 207.6057
[5 empty] | 190.6614 | 191.153
[5 empty] | 195.216 | 196.4977
[5 empty] | 200.4949 | 200.7571
[5 empty] | 192.4979 | 191.8253
[5 empty] | 197.6636 | 197.897
[5 empty] | 195.9205 | 196.5661
[5 empty] | 204.6445 | 203.2538
[5 empty] | 204.2135 | 204.2181
[5 empty] | 203.3534 | 202.0325
[5 empty] | 196.1542 | 196.5526
[5 empty] | 200.9471 | 200.8285
[5 empty] | 198.037 | 199.6229
[5 empty] | 201.9463 | 201.3483
[5 empty] | 196.2482 | 197.2474
[5 empty] | 200.7587 | 199.8529
[5 empty] | 198.3009 | 197.7546
[5 empty] | 202.5675 | 203.5212
[5 empty] | 204.6728 | 204.2605
[5 empty] | 199.8223 | 198.4975
[5 empty] | 191.3737 | 189.1793
[5 empty] | 199.4221 | 199.2189
[5 empty] | 202.3258 | 203.189
[5 empty] | 203.4877 | 203.6978
[5 empty] | 199.0803 | 198.6221
[5 empty] | 200.1273 | 199.6374
[5 empty] | 198.7574 | 199.9938
[5 empty] | 200.4799 | 199.6907
[5 empty] | 201.333 | 202.6908
[5 empty] | 198.9259 | 201.5506
[5 empty] | 200.1382 | 200.0677
[5 empty] | 197.583 | 196.5599
[5 empty] | 197.3891 | 197.9727
[5 empty] | 198.1542 | 197.6355
[5 empty] | 197.0605 | 197.6858
[5 empty] | 198.6067 | 196.3415
[5 empty] | 204.98 | 205.4377
[5 empty] | 195.2549 | 195.9002
[5 empty] | 205.3567 | 207.8841
[5 empty] | 202.4019 | 202.5958
[5 empty] | 203.5312 | 202.6336
[5 empty] | 197.5983 | 198.411
[5 empty] | 202.396 | 201.9636
[5 empty] | 200.6088 | 200.222
[5 empty] | 200.2908 | 200.0162
[5 empty] | 202.2635 | 203.6857
[5 empty] | 208.588 | 207.3093
[5 empty] | 199.3073 | 200.191
[5 empty] | 198.7536 | 196.4849
[5 empty] | 200.6893 | 200.4588
[5 empty] | 206.9976 | 204.583
[5 empty] | 197.2709 | 197.3315
[5 empty] | 199.6377 | 199.0493
[5 empty] | 200.604 | 200.2958
[5 empty] | 206.2373 | 205.4157
[5 empty] | 202.5284 | 202.3881
[5 empty] | 195.5038 | 195.6809
[5 empty] | 197.2935 | 197.3882
[5 empty] | 197.8029 | 198.7071
[5 empty] | 195.9663 | 196.0533
[5 empty] | 201.7511 | 201.9791
[5 empty] | 201.596 | 200.5126
[5 empty] | 196.1103 | 196.0734
[5 empty] | 202.2853 | 202.8766
[5 empty] | 203.6878 | 204.7013
[5 empty] | 198.5787 | 199.1825
[5 empty] | 200.9433 | 199.7779
[5 empty] | 194.4739 | 195.5184
[5 empty] | 200.1324 | 199.8054
[5 empty] | 201.6411 | 201.0592
[5 empty] | 203.815 | 203.9665
[5 empty] | 201.19 | 201.0874
[5 empty] | 200.7143 | 201.537
[5 empty] | 203.5563 | 203.6021
[5 empty] | 198.2494 | 199.0121
[5 empty] | 208.6493 | 207.9997
[5 empty] | 204.9016 | 205.7503
[5 empty] | 197.7391 | 197.553
[5 empty] | 203.9404 | 202.4489
[5 empty] | 198.2439 | 198.6239

## Risk Calculator - Gamma
Global (Average) Risk Calculator - GAMMA Product Distribution
Upper Tolerance | 2 |  |  |  |  | ← Intermediate and plot data in hidden columns here
Measured Product Average | 1 | Upper Tolerance | 2
Measured Prodcut Standard Dev | 0.5 | alpha | 4
Measurement Uncertainty (k=2) ± | 0.75 | beta | 4
Guardband Method | PFA
Manual Acceptance Limit | 2 | Guardband Width | 0.011179
PFA Target | 1% | Upper Acceptance Limit | 1.9888
Acceptance Limit | 1.9888
TUR | 1.3333
Average PFA (No Guardband) | 1.035%
Average PFA(Guardbanded) | 1.002%
Average PFR (Guardbanded) | 7.153%
Worst-case Specific Risk | 48.81%
Process Yield | 95.76%
Process Risk | 4.238%
Process Capability, Cpk | 0.57457
|  | x | TL | TU | AU
|  | -1 | 0 | 2 | 1.9888
|  | 3 | 0 | 2 | 1.9888
|  | N | 1000
|  |  | Product Sample | Measurement Sample
|  |  | 0.78582 | 0.97162
|  |  | 1.4612 | 1.2407
|  |  | 0.69274 | 0.056481
|  |  | 1.3197 | 2.419
|  |  | 0.72098 | 0.81539
|  |  | 1.3226 | 1.4202
|  |  | 1.0521 | 0.71324
|  |  | 0.58269 | 0.80984
|  |  | 0.22051 | 0.48712
|  |  | 0.3662 | 0.51814
|  |  | 1.1588 | 0.92667
|  |  | 1.4432 | 1.2048
|  |  | 0.90423 | 0.93217
|  |  | 1.2393 | 1.231
|  |  | 0.56171 | 0.19871
|  |  | 2.2013 | 1.7249
|  |  | 0.77997 | 0.73413
|  |  | 0.9994 | 0.60583
|  |  | 1.0295 | 1.0988
|  |  | 1.6865 | 1.799
|  |  | 1.4289 | 1.4694
|  |  | 0.53139 | 0.19957
|  |  | 1.3703 | 1.7723
|  |  | 1.5328 | 2.4841
|  |  | 0.81438 | 1.1399
|  |  | 0.40313 | 0.97183
|  |  | 0.97771 | 1.0008
|  |  | 2.4209 | 3.0667
|  |  | 0.67471 | 0.13837
|  |  | 0.73789 | 0.92136
|  |  | 0.51816 | 0.14801
|  |  | 0.77991 | 1.1498
|  |  | 0.63356 | 0.39033
|  |  | 1.0925 | 0.65405
|  |  | 2.7494 | 2.769
|  |  | 0.82873 | 0.56187
|  |  | 1.0856 | 1.1021
|  |  | 0.38791 | -0.15859
|  |  | 0.7531 | 1.0077
|  |  | 1.0292 | 0.13863
|  |  | 0.44546 | 0.4681
|  |  | 0.37896 | 0.075888
|  |  | 0.44566 | 0.73023
|  |  | 0.44909 | 0.47768
|  |  | 1.1084 | 0.62448
|  |  | 0.87452 | -0.15273
|  |  | 0.64753 | 0.047135
|  |  | 0.94946 | 0.14632
|  |  | 1.9259 | 2.1981
|  |  | 1.6171 | 1.7635
|  |  | 0.40137 | -0.086863
|  |  | 1.1314 | 1.5858
|  |  | 0.71099 | 0.55161
|  |  | 1.4629 | 2.0844
|  |  | 0.26646 | 0.36522
|  |  | 1.1786 | 1.4297
|  |  | 1.2529 | 1.5211
|  |  | 1.102 | 0.79273
|  |  | 0.50602 | 0.48045
|  |  | 1.2926 | 0.88033
|  |  | 2.4053 | 2.0451
|  |  | 0.78089 | 1.1637
|  |  | 0.99954 | 0.75901
|  |  | 1.817 | 2.0774
|  |  | 0.51968 | 0.56223
|  |  | 0.31279 | 0.64514
|  |  | 1.5898 | 2.3251
|  |  | 0.44138 | 0.65835
|  |  | 1.0754 | 1.0033
|  |  | 1.8338 | 1.7057
|  |  | 1.2631 | 1.7423
|  |  | 0.62286 | 0.20108
|  |  | 1.1425 | 0.86397
|  |  | 1.4573 | 1.4511
|  |  | 0.47473 | 1.3373
|  |  | 0.79468 | 0.9941
|  |  | 0.49122 | -0.10457
|  |  | 0.47498 | 0.66019
|  |  | 0.60274 | 0.60458
|  |  | 1.1624 | 1.1868
|  |  | 0.5326 | 0.70098
|  |  | 1.4987 | 1.5662
|  |  | 1.8426 | 2.0416
|  |  | 1.2801 | 0.90817
|  |  | 0.58103 | 0.62642
|  |  | 0.78097 | 0.38277
|  |  | 1.4011 | 1.2381
|  |  | 0.78249 | 0.80076
|  |  | 1.2329 | 1.1233
|  |  | 0.63983 | 1.1424
|  |  | 0.22217 | 0.6304
|  |  | 0.27118 | 0.18574
|  |  | 0.92659 | 1.1946
|  |  | 0.64076 | 0.042731
|  |  | 0.89542 | 0.62335
|  |  | 0.9631 | 0.87379
|  |  | 0.66433 | 0.33423
|  |  | 1.289 | 1.7856
|  |  | 0.87242 | 0.57544
|  |  | 0.63463 | 1.4491
|  |  | 0.54993 | -0.25658
|  |  | 1.0047 | 0.95827
|  |  | 0.88592 | 0.70088
|  |  | 0.79004 | 0.81293
|  |  | 1.0705 | 1.279
|  |  | 1.2457 | 1.485
|  |  | 1.0533 | 1.163
|  |  | 1.1272 | 1.676
|  |  | 0.17582 | -0.029374
|  |  | 0.19696 | -0.63221
|  |  | 0.40485 | 0.40956
|  |  | 0.84196 | 0.7182
|  |  | 2.3452 | 2.7501
|  |  | 0.13437 | -0.27722
|  |  | 0.48304 | 0.36228
|  |  | 1.0286 | 1.5139
|  |  | 0.92438 | 1.0793
|  |  | 0.52328 | 0.64181
|  |  | 1.5471 | 1.4617
|  |  | 1.2527 | 1.5195
|  |  | 1.7121 | 0.83609
|  |  | 1.0083 | 1.309
|  |  | 0.30636 | 0.20675
|  |  | 0.75398 | 1.4857
|  |  | 0.49618 | 0.44873
|  |  | 1.3949 | 1.6704
|  |  | 0.48617 | 0.93283
|  |  | 0.7806 | -0.17612
|  |  | 1.3298 | 1.5892
|  |  | 1.0508 | 0.63295
|  |  | 1.1284 | 0.78665
|  |  | 0.75178 | 0.4112
|  |  | 0.23512 | 0.13487
|  |  | 1.5084 | 1.4475
|  |  | 1.4565 | 1.2779
|  |  | 1.6845 | 1.19
|  |  | 0.50135 | 1.1474
|  |  | 0.53803 | 0.6114
|  |  | 0.86177 | 1.6121
|  |  | 0.57616 | 0.36586
|  |  | 1.2117 | 1.2795
|  |  | 1.3772 | 1.5113
|  |  | 1.8651 | 1.8598
|  |  | 0.91295 | -0.091319
|  |  | 0.66684 | 0.81324
|  |  | 0.86416 | 0.65381
|  |  | 0.93521 | 0.77888
|  |  | 0.99308 | 1.4665
|  |  | 1.8755 | 2.1733
|  |  | 1.493 | 2.3802
|  |  | 0.87257 | 1.139
|  |  | 1.7635 | 1.8553
|  |  | 2.0423 | 2.2813
|  |  | 0.21076 | 1.2534
|  |  | 1.495 | 1.2012
|  |  | 0.83791 | 0.41566
|  |  | 1.0443 | 0.95336
|  |  | 1.5649 | 2.4159
|  |  | 1.8891 | 2.2621
|  |  | 2.0121 | 2.5325
|  |  | 1.3704 | 1.5791
|  |  | 0.68068 | 0.64674
|  |  | 1.1642 | 1.1816
|  |  | 1.5173 | 1.6275
|  |  | 0.66774 | 0.66829
|  |  | 1.3232 | 1.9449
|  |  | 0.82016 | 0.55421
|  |  | 0.94966 | 0.29237
|  |  | 1.6558 | 1.4353
|  |  | 1.0622 | 1.1414
|  |  | 0.33619 | 0.94871
|  |  | 0.88534 | 0.81328
|  |  | 0.57101 | 0.63571
|  |  | 1.3508 | 1.1389
|  |  | 2.0455 | 1.2998
|  |  | 1.0985 | 1.8571
|  |  | 0.539 | 0.39248
|  |  | 0.39646 | 0.21291
|  |  | 0.5408 | 1.4082
|  |  | 0.63815 | 1.0189
|  |  | 1.1599 | 1.4869
|  |  | 0.64295 | 0.44075
|  |  | 0.70948 | 1.1969
|  |  | 0.5653 | 0.47469
|  |  | 1.8679 | 1.6366
|  |  | 0.34682 | 0.2765
|  |  | 0.47179 | 0.36887
|  |  | 1.013 | 0.46285
|  |  | 1.7791 | 1.6951
|  |  | 0.42016 | 1.0402
|  |  | 0.87945 | 1.3236
|  |  | 0.6549 | 0.6573
|  |  | 1.9056 | 2.0789
|  |  | 0.58333 | 1.0082
|  |  | 0.59661 | 0.81281
|  |  | 1.6987 | 1.3
|  |  | 1.3636 | 1.5327
|  |  | 1.7271 | 0.76276
|  |  | 0.76576 | 0.81701
|  |  | 1.2139 | 0.86465
|  |  | 1.4695 | 2.1281
|  |  | 1.168 | 1.5644
|  |  | 0.90977 | 0.62379
|  |  | 0.99737 | 1.0142
|  |  | 0.71971 | 0.78046
|  |  | 1.0615 | 0.8894
|  |  | 0.40483 | 0.27734
|  |  | 1.1791 | 1.0603
|  |  | 1.4386 | 1.9107
|  |  | 0.67272 | 0.76156
|  |  | 0.58543 | 0.348
|  |  | 0.72732 | 0.60471
|  |  | 2.1024 | 1.9027
|  |  | 1.4534 | 1.159
|  |  | 0.26771 | -0.099402
|  |  | 1.1924 | 1.147
|  |  | 0.74244 | 1.1127
|  |  | 0.98752 | 1.1971
|  |  | 1.3622 | 1.2841
|  |  | 0.37969 | 0.6169
|  |  | 0.81173 | 0.43417
|  |  | 0.68011 | 0.8449
|  |  | 1.0543 | 1.0993
|  |  | 0.99499 | 1.3537
|  |  | 0.71284 | 0.5259
|  |  | 1.8324 | 1.855
|  |  | 0.82695 | 0.88208
|  |  | 1.6969 | 1.4625
|  |  | 0.50115 | 0.45959
|  |  | 1.034 | 1.223
|  |  | 1.2435 | 0.83499
|  |  | 0.93613 | 0.4455
|  |  | 0.59258 | 1.1945
|  |  | 1.1975 | 0.67692
|  |  | 0.96194 | 0.78621
|  |  | 1.2354 | 1.3624
|  |  | 1.37 | 0.87851
|  |  | 0.61832 | 0.75037
|  |  | 0.78627 | 1.3332
|  |  | 0.72508 | 0.31054
|  |  | 0.51472 | 0.86079
|  |  | 1.4442 | 1.7183
|  |  | 0.39155 | 0.14427
|  |  | 1.381 | 1.3156
|  |  | 1.1682 | 1.7711
|  |  | 0.48369 | 0.74854
|  |  | 1.0982 | 0.86385
|  |  | 0.66233 | 0.50123
|  |  | 0.87247 | 0.74913
|  |  | 0.72705 | 0.58429
|  |  | 1.1013 | 0.74597
|  |  | 0.54532 | 1.132
|  |  | 1.7613 | 1.724
|  |  | 1.5868 | 1.6382
|  |  | 0.76331 | 0.42215
|  |  | 1.1497 | 0.51312
|  |  | 1.0808 | 1.0438
|  |  | 0.44737 | 0.10497
|  |  | 1.1012 | 1.496
|  |  | 1.2378 | 1.5061
|  |  | 0.44293 | 0.27336
|  |  | 0.16305 | 0.52539
|  |  | 0.19301 | -0.41183
|  |  | 0.58138 | 0.98082
|  |  | 0.66203 | 0.58792
|  |  | 1.8949 | 1.7314
|  |  | 0.71109 | 0.59816
|  |  | 3.1434 | 3.2308
|  |  | 2.1058 | 2.079
|  |  | 0.71335 | -0.47435
|  |  | 0.65179 | 0.72431
|  |  | 1.5129 | 1.266
|  |  | 1.2597 | 1.5699
|  |  | 0.57859 | 0.71855
|  |  | 1.1642 | 1.4299
|  |  | 0.61333 | 0.62747
|  |  | 1.003 | 0.87745
|  |  | 1.4441 | 1.8601
|  |  | 0.16448 | 0.43505
|  |  | 1.0974 | 1.1051
|  |  | 0.96008 | 0.67529
|  |  | 0.75455 | 0.94498
|  |  | 0.56521 | 1.4045
|  |  | 0.59277 | 0.84489
|  |  | 0.5041 | 0.36036
|  |  | 0.64117 | -0.043649
|  |  | 0.39956 | 0.77837
|  |  | 0.94309 | 1.2387
|  |  | 0.76536 | 0.26415
|  |  | 0.26676 | -0.32029
|  |  | 1.0782 | 0.80776
|  |  | 2.1679 | 1.7277
|  |  | 1.1576 | 0.6389
|  |  | 1.1432 | 1.1246
|  |  | 1.208 | 1.6989
|  |  | 1.0441 | 1.7221
|  |  | 2.0971 | 2.2038
|  |  | 1.0992 | 0.33579
|  |  | 0.23411 | 0.39272
|  |  | 0.92742 | 1.0887
|  |  | 0.3728 | -0.17782
|  |  | 1.4717 | 0.82337
|  |  | 1.5317 | 1.7532
|  |  | 0.90519 | 0.97539
|  |  | 1.5465 | 1.1857
|  |  | 0.72047 | 0.59449
|  |  | 0.37254 | 0.64948
|  |  | 0.54684 | 0.50492
|  |  | 1.6757 | 1.0543
|  |  | 2.395 | 2.1782
|  |  | 0.8361 | 0.90534
|  |  | 1.1891 | 0.6458
|  |  | 0.93886 | 1.3912
|  |  | 1.3222 | 0.74849
|  |  | 1.0533 | 1.4054
|  |  | 0.36761 | 0.19304
|  |  | 0.74245 | 0.36764
|  |  | 0.90395 | 0.61294
|  |  | 1.4998 | 2.0651
|  |  | 1.2394 | 1.7949
|  |  | 0.45305 | 0.4834
|  |  | 1.163 | 0.9416
|  |  | 1.4743 | 1.4374
|  |  | 0.92871 | 1.4006
|  |  | 1.2976 | 1.3832
|  |  | 0.55395 | 1.0145
|  |  | 0.94574 | 0.82447
|  |  | 1.3394 | 0.72641
|  |  | 1.0962 | 1.0547
|  |  | 0.79378 | 0.12058
|  |  | 1.053 | 0.79899
|  |  | 1.0528 | 1.184
|  |  | 0.59525 | 0.20093
|  |  | 1.286 | 1.2496
|  |  | 0.49773 | 0.6338
|  |  | 0.94502 | 1.0331
|  |  | 1.7335 | 2.2904
|  |  | 1.3501 | 1.0044
|  |  | 0.64622 | 0.46985
|  |  | 1.0767 | 0.98309
|  |  | 1.1955 | 1.2571
|  |  | 0.87437 | 1.6481
|  |  | 0.64066 | 0.7413
|  |  | 1.111 | 1.1026
|  |  | 0.37993 | 0.77017
|  |  | 0.92781 | 0.7794
|  |  | 0.86553 | 1.3522
|  |  | 1.0886 | 0.80003
|  |  | 0.64674 | 1.165
|  |  | 0.68475 | 0.67082
|  |  | 1.4579 | 1.3325
|  |  | 0.067402 | 0.39716
|  |  | 0.31297 | -0.15561
|  |  | 0.63689 | 0.32247
|  |  | 0.64903 | 1.0071
|  |  | 0.87592 | 0.35236
|  |  | 1.476 | 1.4064
|  |  | 0.79705 | 1.3554
|  |  | 1.1972 | 0.58184
|  |  | 0.50324 | 1.1049
|  |  | 0.6357 | 0.52675
|  |  | 0.72095 | 0.85053
|  |  | 1.7779 | 1.0537
|  |  | 0.92079 | 0.8025
|  |  | 0.7651 | 0.52465
|  |  | 0.41007 | 0.46018
|  |  | 1.2626 | 0.77755
|  |  | 0.43668 | 0.7516
|  |  | 0.96446 | 0.53235
|  |  | 0.953 | 1.3208
|  |  | 1.6163 | 2.1996
|  |  | 1.3968 | 1.1577
|  |  | 0.39138 | 0.25374
|  |  | 0.94529 | 0.41101
|  |  | 1.0369 | 1.3438
|  |  | 0.2495 | 0.0074172
|  |  | 1.3993 | 1.4929
|  |  | 0.71827 | 0.39041
|  |  | 0.94374 | 1.0955
|  |  | 1.368 | 1.3283
|  |  | 0.9298 | 1.1465
|  |  | 1.5526 | 0.35946
|  |  | 1.4872 | 1.8137
|  |  | 1.3297 | 0.97105
|  |  | 1.9595 | 2.1293
|  |  | 0.87438 | 1.0332
|  |  | 1.703 | 1.58
|  |  | 0.3352 | -1.0169
|  |  | 0.7783 | 0.90408
|  |  | 0.40211 | 0.44949
|  |  | 1.0588 | 1.0651
|  |  | 1.0207 | 0.9836
|  |  | 1.3619 | 0.90163
|  |  | 1.7052 | 1.4299
|  |  | 0.39488 | 0.22596
|  |  | 1.0203 | 1.4255
|  |  | 1.6057 | 1.9131
|  |  | 0.065472 | 0.1841
|  |  | 1.0291 | 0.99517
|  |  | 0.66502 | 0.60791
|  |  | 1.7919 | 1.7615
|  |  | 1.2497 | 1.3581
|  |  | 1.723 | 1.6002
|  |  | 2.3608 | 2.319
|  |  | 1.3 | 1.0491
|  |  | 0.51401 | 0.043105
|  |  | 0.69859 | 0.57556
|  |  | 1.9748 | 2.3194
|  |  | 1.8266 | 1.6819
|  |  | 0.33864 | 0.70357
|  |  | 1.4892 | 1.9781
|  |  | 0.57005 | 0.5509
|  |  | 0.78094 | 0.40647
|  |  | 0.71615 | 1.0718
|  |  | 2.1104 | 2.2551
|  |  | 0.75628 | 0.81229
|  |  | 0.88389 | 0.3383
|  |  | 0.31409 | 0.062952
|  |  | 1.3505 | 1.6421
|  |  | 0.69645 | 0.85989
|  |  | 2.2282 | 1.9598
|  |  | 0.51444 | 0.62346
|  |  | 1.0311 | 0.97059
|  |  | 1.0837 | 0.82909
|  |  | 1.1999 | 1.2542
|  |  | 0.60539 | 0.38072
|  |  | 0.65959 | 1.2763
|  |  | 0.85502 | 1.061
|  |  | 1.1708 | 1.3383
|  |  | 0.59084 | 0.53672
|  |  | 2.779 | 2.6669
|  |  | 1.4139 | 1.3911
|  |  | 0.53355 | 0.39263
|  |  | 2.0384 | 2.1216
|  |  | 0.51294 | 0.47051
|  |  | 1.0714 | 0.74954
|  |  | 0.82602 | 0.64333
|  |  | 0.70184 | 0.91777
|  |  | 1.1725 | 0.5732
|  |  | 0.40272 | 0.15337
|  |  | 0.93166 | 1.0039
|  |  | 1.3908 | 1.0937
|  |  | 0.2929 | -0.12477
|  |  | 0.65825 | 0.80658
|  |  | 1.3038 | 1.1061
|  |  | 0.62148 | 0.77201
|  |  | 0.38925 | 0.37513
|  |  | 0.83165 | 0.64284
|  |  | 1.0946 | 1.0305
|  |  | 1.1453 | 1.1992
|  |  | 0.90242 | 0.84746
|  |  | 1.0084 | 0.61536
|  |  | 1.285 | 1.9504
|  |  | 1.0307 | 0.39611
|  |  | 1.1587 | 1.098
|  |  | 2.6478 | 1.9701
|  |  | 0.37959 | 0.017539
|  |  | 1.6591 | 1.1125
|  |  | 0.70287 | 0.35218
|  |  | 0.52128 | 0.13836
|  |  | 0.50069 | 0.20595
|  |  | 0.86486 | 0.90875
|  |  | 0.37797 | 0.57422
|  |  | 1.6761 | 2.351
|  |  | 1.3162 | 1.0611
|  |  | 0.15321 | 0.2336
|  |  | 0.46882 | 0.20914
|  |  | 1.9822 | 2.4428
|  |  | 1.6489 | 1.6415
|  |  | 0.87231 | 0.95019
|  |  | 0.60331 | 0.18008
|  |  | 1.5053 | 1.7637
|  |  | 1.0836 | 0.56143
|  |  | 1.2414 | 1.7198
|  |  | 0.82257 | 0.49999
|  |  | 0.80657 | 0.88741
|  |  | 0.94424 | 1.6519
|  |  | 1.1643 | 1.5949
|  |  | 0.90499 | 1.334
|  |  | 0.28703 | 0.58922
|  |  | 0.66124 | 0.032827
|  |  | 0.63219 | 0.83079
|  |  | 1.3384 | 1.3319
|  |  | 0.70142 | 0.92593
|  |  | 1.4738 | 1.7909
|  |  | 1.4322 | 1.7762
|  |  | 0.51885 | 1.0378
|  |  | 0.75825 | 0.80189
|  |  | 1.1418 | 0.94918
|  |  | 0.60337 | 0.14989
|  |  | 1.4079 | 1.4488
|  |  | 1.5487 | 1.094
|  |  | 0.74128 | 0.46163
|  |  | 0.8092 | 0.132
|  |  | 1.55 | 1.3538
|  |  | 0.42655 | 0.48902
|  |  | 0.84078 | 1.0878
|  |  | 0.65602 | 1.203
|  |  | 1.0888 | 1.5823
|  |  | 1.3618 | 1.1797
|  |  | 1.8105 | 2.1893
|  |  | 0.27844 | 0.89628
|  |  | 0.41972 | 0.32866
|  |  | 1.5553 | 1.2393
|  |  | 0.59398 | -0.2293
|  |  | 0.92226 | 1.4218
|  |  | 1.8805 | 2.2036
|  |  | 0.47042 | 0.6189
|  |  | 0.18945 | 0.46363
|  |  | 0.39033 | -0.34828
|  |  | 0.94394 | 0.69385
|  |  | 1.0349 | 1.2374
|  |  | 1.4319 | 1.6667
|  |  | 0.99805 | 0.77882
|  |  | 0.80395 | 1.5168
|  |  | 1.8072 | 1.8083
|  |  | 1.8998 | 1.1504
|  |  | 1.3418 | 1.3496
|  |  | 1.148 | 1.3677
|  |  | 0.71395 | 1.1662
|  |  | 1.7815 | 2.0123
|  |  | 1.3092 | 1.0134
|  |  | 0.8884 | 0.82492
|  |  | 0.69016 | 0.47405
|  |  | 0.84476 | 0.85369
|  |  | 0.89075 | 0.49806
|  |  | 0.81089 | 1.4208
|  |  | 0.48976 | 1.5401
|  |  | 2.2569 | 1.8599
|  |  | 0.53851 | 0.61129
|  |  | 0.52762 | 0.89884
|  |  | 1.4822 | 1.0297
|  |  | 0.49638 | 0.66785
|  |  | 1.6021 | 1.337
|  |  | 0.5732 | 0.40006
|  |  | 1.9684 | 2.1542
|  |  | 1.4414 | 0.52526
|  |  | 0.71279 | 0.97922
|  |  | 0.49623 | 0.52883
|  |  | 0.68375 | 0.52067
|  |  | 0.7593 | 0.68133
|  |  | 0.75773 | -0.10104
|  |  | 1.7493 | 1.6741
|  |  | 0.89353 | 0.65856
|  |  | 0.61452 | 1.2196
|  |  | 1.1232 | 0.53148
|  |  | 1.4833 | 1.1064
|  |  | 0.86848 | 0.71264
|  |  | 0.74083 | 0.94888
|  |  | 0.72489 | 0.76096
|  |  | 0.76317 | 1.0579
|  |  | 0.9463 | 0.57533
|  |  | 0.46583 | 0.2844
|  |  | 0.96529 | 1.1335
|  |  | 0.19342 | 0.3075
|  |  | 1.0512 | 1.4197
|  |  | 0.98915 | 0.40558
|  |  | 0.50058 | 0.50808
|  |  | 0.72044 | 0.84905
|  |  | 2.3462 | 2.2817
|  |  | 1.7122 | 1.9052
|  |  | 0.45123 | -0.4776
|  |  | 0.71166 | 0.40123
|  |  | 0.86179 | 0.50263
|  |  | 0.97133 | 0.85247
|  |  | 1.0052 | 1.0182
|  |  | 2.0763 | 2.2398
|  |  | 0.97538 | 0.94328
|  |  | 1.7153 | 1.9428
|  |  | 1.5301 | 1.691
|  |  | 1.5185 | 1.3539
|  |  | 0.65036 | 0.58975
|  |  | 0.97394 | 1.1764
|  |  | 0.31357 | 0.59279
|  |  | 0.51007 | 0.49629
|  |  | 1.2849 | 1.205
|  |  | 0.81026 | 0.9447
|  |  | 0.89889 | 0.85367
|  |  | 1.0589 | 1.1671
|  |  | 1.5848 | 1.6337
|  |  | 0.77596 | 0.82635
|  |  | 0.56324 | 0.44385
|  |  | 1.9345 | 1.673
|  |  | 0.93853 | 1.6104
|  |  | 1.0709 | 0.61119
|  |  | 1.0376 | 1.2945
|  |  | 0.56619 | 1.0723
|  |  | 0.66916 | 0.86746
|  |  | 0.54132 | 0.23105
|  |  | 2.328 | 2.7825
|  |  | 1.2202 | 0.95493
|  |  | 1.8303 | 1.3267
|  |  | 1.384 | 1.7112
|  |  | 0.7524 | 0.74522
|  |  | 1.2653 | 0.79258
|  |  | 1.1158 | 1.1082
|  |  | 0.87111 | 0.74105
|  |  | 1.0046 | 0.82529
|  |  | 0.48671 | 0.51417
|  |  | 1.0601 | 0.71481
|  |  | 1.0104 | 1.0754
|  |  | 0.80167 | 1.0259
|  |  | 1.0203 | 0.53945
|  |  | 1.483 | 1.4886
|  |  | 2.3036 | 3.2029
|  |  | 1.1366 | 1.2807
|  |  | 3.3389 | 3.2804
|  |  | 0.49588 | 1.0514
|  |  | 1.4297 | 1.6995
|  |  | 0.98176 | 0.7725
|  |  | 1.7731 | 2.0278
|  |  | 0.8341 | 0.94816
|  |  | 1.7478 | 1.6541
|  |  | 1.3529 | 1.1739
|  |  | 1.2361 | 1.4015
|  |  | 2.3087 | 2.3473
|  |  | 0.68945 | 0.12805
|  |  | 1.7102 | 1.4912
|  |  | 1.2811 | 1.495
|  |  | 1.4933 | 1.5808
|  |  | 0.70885 | 0.73151
|  |  | 1.8228 | 2.1495
|  |  | 0.21286 | 0.65923
|  |  | 0.92919 | 1.3392
|  |  | 0.4469 | 1.0128
|  |  | 0.3666 | 0.34836
|  |  | 1.3467 | 1.2491
|  |  | 0.4519 | 1.1712
|  |  | 0.72549 | 1.1051
|  |  | 0.77239 | 0.28698
|  |  | 0.9139 | 1.5894
|  |  | 0.92371 | 1.0709
|  |  | 0.49351 | 0.14995
|  |  | 2.5491 | 3.2316
|  |  | 0.29612 | -0.04446
|  |  | 0.70696 | 0.66505
|  |  | 1.2375 | 1.2424
|  |  | 1.404 | 1.3424
|  |  | 2.3433 | 1.955
|  |  | 1.2509 | 1.4843
|  |  | 0.84877 | 0.78316
|  |  | 1.2661 | 1.2851
|  |  | 1.362 | 1.4225
|  |  | 0.6619 | 1.0433
|  |  | 0.49311 | 0.50477
|  |  | 0.78772 | 1.0306
|  |  | 0.58276 | 0.23361
|  |  | 1.393 | 1.3515
|  |  | 0.41988 | 0.7431
|  |  | 0.59136 | 0.055302
|  |  | 1.232 | 1.2411
|  |  | 0.56329 | 0.9198
|  |  | 0.47508 | 0.72882
|  |  | 0.95641 | 0.77581
|  |  | 2.1199 | 1.9497
|  |  | 1.0332 | 1.4722
|  |  | 0.69474 | 1.1234
|  |  | 1.1369 | 1.1733
|  |  | 0.47869 | -0.19242
|  |  | 0.5634 | 0.26749
|  |  | 1.9174 | 1.8631
|  |  | 1.86 | 2.1565
|  |  | 0.95867 | 1.0129
|  |  | 0.4036 | 0.44632
|  |  | 1.2629 | 2.1515
|  |  | 0.58383 | 1.226
|  |  | 1.9399 | 2.0536
|  |  | 0.49111 | 1.0988
|  |  | 0.79828 | 0.48127
|  |  | 0.75768 | 0.19186
|  |  | 1.427 | 1.4986
|  |  | 0.82803 | 0.89801
|  |  | 1.2306 | 1.3709
|  |  | 1.1409 | 1.0092
|  |  | 1.6673 | 0.84836
|  |  | 0.88349 | 0.95238
|  |  | 1.4224 | 1.3847
|  |  | 0.36452 | 0.34663
|  |  | 0.72437 | 1.1435
|  |  | 0.74975 | 0.55291
|  |  | 1.8408 | 0.9349
|  |  | 0.51076 | 0.46418
|  |  | 0.54408 | 0.051782
|  |  | 1.0422 | 1.1636
|  |  | 0.59171 | 1.2082
|  |  | 1.2717 | 1.2547
|  |  | 0.42847 | 0.32991
|  |  | 0.61971 | 0.70917
|  |  | 1.9988 | 1.7351
|  |  | 0.64038 | 0.39408
|  |  | 1.0223 | 1.7123
|  |  | 1.4101 | 1.8548
|  |  | 0.36695 | 1.0192
|  |  | 1.23 | 1.563
|  |  | 0.64627 | 1.4414
|  |  | 1.8117 | 1.4784
|  |  | 0.4063 | -0.022609
|  |  | 0.48587 | 0.57677
|  |  | 0.28707 | 0.20629
|  |  | 0.62554 | 0.16503
|  |  | 0.16059 | 0.62225
|  |  | 0.69805 | 0.84523
|  |  | 0.94749 | 0.98436
|  |  | 1.3236 | 1.4062
|  |  | 0.50133 | 0.5651
|  |  | 0.87513 | 1.0479
|  |  | 0.85839 | 0.21635
|  |  | 1.7368 | 1.6315
|  |  | 0.48566 | 0.3504
|  |  | 1.1594 | 1.3625
|  |  | 1.7715 | 1.6578
|  |  | 0.56417 | -0.019357
|  |  | 1.5589 | 1.9996
|  |  | 1.0889 | 1.1762
|  |  | 1.0403 | 1.0708
|  |  | 0.86967 | 0.93379
|  |  | 0.8241 | 0.52343
|  |  | 1.6712 | 1.2797
|  |  | 0.45827 | 0.6991
|  |  | 0.95913 | 1.0334
|  |  | 0.71446 | 0.040992
|  |  | 0.51751 | 0.545
|  |  | 1.3172 | 1.106
|  |  | 0.66605 | 0.78064
|  |  | 0.30296 | 0.92303
|  |  | 1.2414 | 1.2204
|  |  | 0.18158 | 0.4036
|  |  | 2.6951 | 2.5248
|  |  | 0.65782 | 0.64788
|  |  | 1.6836 | 1.2266
|  |  | 0.63157 | 0.48586
|  |  | 1.1489 | 1.0992
|  |  | 0.761 | 0.27386
|  |  | 0.54941 | 0.083126
|  |  | 0.9117 | 1.2035
|  |  | 2.187 | 1.8703
|  |  | 0.74054 | 0.52473
|  |  | 1.2286 | 1.3452
|  |  | 1.0994 | 1.4503
|  |  | 0.98874 | 1.3259
|  |  | 0.79108 | 0.8108
|  |  | 2.7149 | 2.6198
|  |  | 0.75191 | 0.39432
|  |  | 0.47053 | 0.31516
|  |  | 0.86112 | 0.42262
|  |  | 0.87073 | 0.86815
|  |  | 0.70317 | -0.32626
|  |  | 0.77232 | 0.70014
|  |  | 0.30541 | 0.76059
|  |  | 0.24637 | 0.35671
|  |  | 1.5608 | 1.0175
|  |  | 1.1213 | 0.81631
|  |  | 1.3749 | 1.607
|  |  | 2.0928 | 2.6288
|  |  | 2.7686 | 2.7414
|  |  | 1.3382 | 1.5327
|  |  | 1.1937 | 0.95489
|  |  | 0.72971 | 1.2247
|  |  | 0.40539 | 0.37461
|  |  | 0.55537 | 0.62913
|  |  | 0.69076 | 0.51607
|  |  | 0.29968 | 0.71615
|  |  | 0.92814 | 0.34525
|  |  | 0.43853 | 0.64044
|  |  | 0.31269 | 0.65756
|  |  | 0.52778 | 0.31121
|  |  | 0.58111 | 0.84743
|  |  | 1.3088 | 1.118
|  |  | 1.1338 | 1.3072
|  |  | 1.0963 | 0.79089
|  |  | 0.55033 | 0.12512
|  |  | 0.90503 | 0.66273
|  |  | 0.57884 | 1.1784
|  |  | 0.40646 | 0.50253
|  |  | 1.2588 | 0.97352
|  |  | 0.74201 | 0.70726
|  |  | 1.2571 | 1.1501
|  |  | 0.44828 | -0.096208
|  |  | 0.53326 | 0.80694
|  |  | 1.1075 | 1.239
|  |  | 0.75607 | 0.060822
|  |  | 1.5324 | 1.5444
|  |  | 1.464 | 1.7871
|  |  | 0.59429 | 0.41329
|  |  | 0.32484 | 0.065606
|  |  | 1.0691 | 1.1592
|  |  | 0.35596 | 1.0069
|  |  | 0.5104 | 0.1889
|  |  | 0.72611 | 0.61283
|  |  | 0.88412 | 1.0564
|  |  | 0.89743 | 1.0671
|  |  | 1.5651 | 1.639
|  |  | 1.0366 | 1.2138
|  |  | 0.62333 | 0.34978
|  |  | 0.53058 | 0.55123
|  |  | 0.43742 | 0.5035
|  |  | 2.3459 | 2.0824
|  |  | 1.7608 | 0.97357
|  |  | 0.27054 | 0.90702
|  |  | 1.21 | 1.2807
|  |  | 1.1717 | 1.2647
|  |  | 2.4463 | 2.4443
|  |  | 0.55153 | 0.40933
|  |  | 0.96775 | 0.3043
|  |  | 2.0989 | 1.5336
|  |  | 0.56805 | 0.38335
|  |  | 2.2491 | 2.2389
|  |  | 1.2572 | 0.74175
|  |  | 0.7364 | 0.50664
|  |  | 1.0487 | 1.3241
|  |  | 0.77759 | 0.98019
|  |  | 0.63539 | 0.738
|  |  | 0.66777 | 0.84128
|  |  | 0.74429 | 0.72811
|  |  | 1.502 | 0.9119
|  |  | 0.90353 | 1.2289
|  |  | 0.97845 | 0.9525
|  |  | 0.59921 | 1.0422
|  |  | 0.78034 | 1.1389
|  |  | 1.2781 | 0.89533
|  |  | 0.77233 | 0.55012
|  |  | 1.1774 | 1.2357
|  |  | 0.27176 | 0.33865
|  |  | 1.8606 | 2.1511
|  |  | 1.1421 | 1.0768
|  |  | 1.5014 | 2.053
|  |  | 0.26649 | -0.053402
|  |  | 0.75887 | 0.41918
|  |  | 0.68057 | 1.2179
|  |  | 0.54966 | 0.17743
|  |  | 1.1392 | 1.1596
|  |  | 0.55204 | 0.47028
|  |  | 1.515 | 1.4835
|  |  | 0.61944 | 0.71539
|  |  | 1.7036 | 1.1455
|  |  | 1.2875 | 1.3822
|  |  | 0.26891 | 0.25925
|  |  | 2.1575 | 2.6154
|  |  | 1.0416 | 0.59906
|  |  | 0.98838 | 1.7276
|  |  | 0.60821 | 0.76511
|  |  | 1.1793 | 1.3132
|  |  | 0.50256 | 0.78883
|  |  | 0.95069 | 1.2848
|  |  | 0.57466 | 0.44384
|  |  | 0.91545 | 0.5686
|  |  | 0.88034 | 1.532
|  |  | 1.1396 | 1.2104
|  |  | 1.1134 | 1.4397
|  |  | 0.87813 | 0.72709
|  |  | 0.20321 | 0.13806
|  |  | 1.0822 | 1.3107
|  |  | 0.83647 | 0.52649
|  |  | 1.0961 | 1.2116
|  |  | 1.3177 | 1.2783
|  |  | 0.87009 | 0.55395
|  |  | 1.2321 | 1.1229
|  |  | 0.39486 | 0.49263
|  |  | 0.58784 | 0.46155
|  |  | 0.56847 | 0.94171
|  |  | 0.64802 | 0.29677
|  |  | 1.7668 | 1.8152
|  |  | 0.49571 | 0.19827
|  |  | 0.72898 | 0.19
|  |  | 1.5762 | 1.6113
|  |  | 0.80729 | 0.9601
|  |  | 0.33997 | 0.31305
|  |  | 0.92237 | 1.116
|  |  | 0.75279 | 1.1865
|  |  | 0.43498 | 0.72064
|  |  | 0.98524 | 1.7248
|  |  | 1.5264 | 1.5705
|  |  | 0.89286 | 0.67501
|  |  | 0.61328 | 0.65411
|  |  | 0.93759 | 0.82741
|  |  | 0.76025 | 0.83834
|  |  | 0.79646 | 1.4117
|  |  | 1.212 | 1.3859
|  |  | 1.4946 | 1.52
|  |  | 1.0869 | 1.005
|  |  | 0.31709 | 0.28659
|  |  | 1.2662 | 1.4484
|  |  | 0.73666 | 0.78768
|  |  | 1.6827 | 1.6031
|  |  | 1.1892 | 1.2272
|  |  | 0.3488 | 0.79877
|  |  | 1.8369 | 1.6631
|  |  | 0.83216 | 0.49846
|  |  | 0.97623 | 1.0824
|  |  | 1.3944 | 1.1873
|  |  | 1.0282 | 1.3877
|  |  | 0.78608 | 0.69634
|  |  | 0.63867 | 0.68214
|  |  | 1.3098 | 1.2245
|  |  | 1.4724 | 0.97664
|  |  | 1.5915 | 2.0544
|  |  | 1.1043 | 0.71725
|  |  | 0.3514 | 0.82904
|  |  | 0.627 | 0.90181
|  |  | 0.83698 | 0.57918
|  |  | 1.022 | 1.3674
|  |  | 0.42812 | 1.0122
|  |  | 0.63466 | 0.79161
|  |  | 0.96216 | 0.86225
|  |  | 1.4602 | 2.0366
|  |  | 0.52187 | 0.53704
|  |  | 1.4692 | 1.9175
|  |  | 0.73679 | 0.7232
|  |  | 1.4675 | 1.9267
|  |  | 1.0328 | 0.44975
|  |  | 0.70664 | -0.015259
|  |  | 1.2759 | 1.0299
|  |  | 0.46572 | 1.5199
|  |  | 1.4042 | 1.069
|  |  | 1.7394 | 2.2012
|  |  | 0.85104 | 0.5943
|  |  | 1.6384 | 2.0437
|  |  | 0.87458 | 1.2034
|  |  | 1.0963 | 0.89154
|  |  | 1.6603 | 1.3917
|  |  | 0.7359 | 0.8986
|  |  | 1.3165 | 2.0569
|  |  | 0.89286 | -0.16169
|  |  | 0.40858 | -0.048863
|  |  | 0.53898 | 0.43583
|  |  | 1.6055 | 1.9276
|  |  | 0.93692 | 1.229
|  |  | 0.91265 | 1.1584
|  |  | 0.57239 | 0.52396
|  |  | 1.2422 | 1.8262
|  |  | 0.44884 | 1.3092
|  |  | 1.2696 | 1.231
|  |  | 0.41692 | 0.38652
|  |  | 1.0762 | 0.73432
|  |  | 0.95621 | 0.69024
|  |  | 1.2665 | 0.88531
|  |  | 1.2027 | 0.73758
|  |  | 1.176 | 0.61531
|  |  | 0.3072 | 0.38365
|  |  | 2.5994 | 2.4249
|  |  | 1.0463 | 0.78259
|  |  | 0.51151 | 0.18911
|  |  | 1.4429 | 1.7926
|  |  | 0.58116 | 0.93895
|  |  | 2.1227 | 2.6408
|  |  | 1.426 | 0.91274
|  |  | 1.2868 | 0.71908
|  |  | 0.75165 | 0.9652
|  |  | 1.8908 | 2.0089
|  |  | 1.0296 | 1.2356
|  |  | 0.47144 | 0.086539
|  |  | 0.95555 | 0.37785
|  |  | 0.33924 | 0.22415
|  |  | 0.97487 | 0.54103
|  |  | 0.84267 | 0.72151
|  |  | 0.30057 | 0.062839
|  |  | 0.44618 | 0.57495
|  |  | 0.80521 | 0.72289
|  |  | 1.7246 | 1.3886
|  |  | 1.6495 | 1.7768
|  |  | 0.6299 | 1.0057
|  |  | 0.81856 | 0.36747
|  |  | 1.4786 | 1.6204
|  |  | 0.85878 | 1.0565
|  |  | 0.50426 | 0.41922
|  |  | 0.5728 | 1.1739
|  |  | 1.6619 | 1.2745
|  |  | 1.2686 | 1.5588
|  |  | 0.7412 | 0.54567
|  |  | 0.31995 | 0.34616
|  |  | 0.28636 | 0.073873
|  |  | 1.1527 | 1.4516
|  |  | 0.91196 | 0.8673
|  |  | 0.96577 | 0.44027
|  |  | 1.432 | 1.3768
|  |  | 0.5043 | 0.54966
|  |  | 0.38963 | -0.13231
|  |  | 1.4572 | 0.90571
|  |  | 1.2858 | 1.0164
|  |  | 1.5385 | 1.9727
|  |  | 0.70584 | 0.3004
|  |  | 0.94742 | 1.2187
|  |  | 0.31201 | 0.028272
|  |  | 1.4741 | 1.7759
|  |  | 0.60706 | 0.42749
|  |  | 0.82067 | 0.64319
|  |  | 0.65439 | 0.55919
|  |  | 1.9486 | 1.8641
|  |  | 0.94052 | 0.80856
|  |  | 0.84423 | 0.90997
|  |  | 0.90876 | 1.5211
|  |  | 0.29155 | 0.83209
|  |  | 1.1406 | 0.89658
|  |  | 0.76362 | 0.65705
|  |  | 1.0258 | 0.7499
|  |  | 0.84313 | 0.5187
|  |  | 0.65422 | 0.47747
|  |  | 0.74549 | 0.5985
|  |  | 0.61996 | 0.34773
|  |  | 0.52076 | 0.29715

## Specific Risk
Plot Probability of Conformance for the measurement uncertainty and tolerance
Specific Risk Calculator | [9 empty] | ← Intermediate and plot data in hidden columns here
Tolerance | 0 | ± | 8 | Lower Tolerance | -8 | Measured | Specific Risk, Lower | Specific Risk, Upper Limit | Probability of Conformance
Measurement Uncertainty (k=2) ± | 3 |  |  | Upper Tolerance | 8 | -16 | 1 | 0 | 4.8213e-08
Minimum Probability of Conformance | 75% |  |  |  |  | -15.6768 | 1 | 0 | 1.5452e-07
|  |  |  | Range Lower | -8 | -15.3535 | 1 | 0 | 4.7347e-07
Lower Acceptance Limit | -6.9883 |  |  | Range Upper | 8 | -15.0303 | 1 | 0 | 1.3871e-06
Upper Acceptance Limit | 6.9883 |  |  |  |  | -14.7071 | 1 | 0 | 3.8858e-06
[6 empty] | -14.3838 | 0.99999 | 0 | 1.0411e-05
|  |  |  | Range | 16 | -14.0606 | 0.99997 | 0 | 2.668e-05
|  |  |  | Low | -16 | -13.7374 | 0.99993 | 0 | 6.5408e-05
|  |  |  | High | 16 | -13.4141 | 0.99985 | 0 | 0.00015344
[6 empty] | -13.0909 | 0.99966 | 0 | 0.00034447
|  |  |  | Lower Guardband Width | 1.0117 | -12.7677 | 0.99926 | 0 | 0.00074032
|  |  |  | Upper Guardband Width | 1.0117 | -12.4444 | 0.99848 | 0 | 0.0015235
|  |  |  | Lower Acceptance Limit | -6.9883 | -12.1212 | 0.997 | 0 | 0.0030028
|  |  |  | Upper Acceptance Limit | 6.9883 | -11.798 | 0.99433 | 0 | 0.0056709
[6 empty] | -11.4747 | 0.98973 | 0 | 0.010266
[6 empty] | -11.1515 | 0.98218 | 0 | 0.01782
|  |  |  | Plot Tolerances: |  | -10.8283 | 0.97032 | 0 | 0.02968
|  |  |  | 0 | -8 | -10.5051 | 0.95254 | 0 | 0.047456
|  |  |  | 1.2 | -8 | -10.1818 | 0.9271 | 0 | 0.072898
[6 empty] | -9.8586 | 0.89234 | 0 | 0.10766
|  |  |  | 0 | 8 | -9.5354 | 0.84698 | 0 | 0.15302
|  |  |  | 1.2 | 8 | -9.2121 | 0.79048 | 0 | 0.20952
[6 empty] | -8.8889 | 0.72327 | 0 | 0.27673
|  |  |  | 0 | -6.9883 | -8.5657 | 0.64695 | 0 | 0.35305
|  |  |  | 1.2 | -6.9883 | -8.2424 | 0.5642 | 0 | 0.4358
[6 empty] | -7.9192 | 0.47852 | 0 | 0.52148
|  |  |  | 0 | 6.9883 | -7.596 | 0.39383 | 0 | 0.60617
|  |  |  | 1.2 | 6.9883 | -7.2727 | 0.31389 | 0 | 0.68611
[6 empty] | -6.9495 | 0.24186 | 0 | 0.75814
[6 empty] | -6.6263 | 0.17988 | 0 | 0.82012
[6 empty] | -6.303 | 0.12896 | 0 | 0.87104
[6 empty] | -5.9798 | 0.089022 | 0 | 0.91098
[6 empty] | -5.6566 | 0.05911 | 0 | 0.94089
[6 empty] | -5.3333 | 0.03772 | 0 | 0.96228
[6 empty] | -5.0101 | 0.023116 | 0 | 0.97688
[6 empty] | -4.6869 | 0.013596 | 0 | 0.9864
[6 empty] | -4.3636 | 0.0076702 | 0 | 0.99233
[6 empty] | -4.0404 | 0.0041486 | 0 | 0.99585
[6 empty] | -3.7172 | 0.0021504 | 2.7756e-15 | 0.99785
[6 empty] | -3.3939 | 0.0010678 | 1.5321e-14 | 0.99893
[6 empty] | -3.0707 | 0.00050775 | 7.8826e-14 | 0.99949
[6 empty] | -2.7475 | 0.00023116 | 3.8902e-13 | 0.99977
[6 empty] | -2.4242 | 0.00010073 | 1.833e-12 | 0.9999
[6 empty] | -2.101 | 4.2006e-05 | 8.2527e-12 | 0.99996
[6 empty] | -1.7778 | 1.6759e-05 | 3.5503e-11 | 0.99998
[6 empty] | -1.4545 | 6.3959e-06 | 1.4594e-10 | 0.99999
[6 empty] | -1.1313 | 2.3346e-06 | 5.7329e-10 | 1
[6 empty] | -0.80808 | 8.1495e-07 | 2.1521e-09 | 1
[6 empty] | -0.48485 | 2.7201e-07 | 7.7216e-09 | 1
[6 empty] | -0.16162 | 8.6801e-08 | 2.648e-08 | 1
[6 empty] | 0.16162 | 2.648e-08 | 8.6801e-08 | 1
[6 empty] | 0.48485 | 7.7216e-09 | 2.7201e-07 | 1
[6 empty] | 0.80808 | 2.1521e-09 | 8.1495e-07 | 1
[6 empty] | 1.1313 | 5.7329e-10 | 2.3346e-06 | 1
[6 empty] | 1.4545 | 1.4594e-10 | 6.3959e-06 | 0.99999
[6 empty] | 1.7778 | 3.5503e-11 | 1.6759e-05 | 0.99998
[6 empty] | 2.101 | 8.2527e-12 | 4.2006e-05 | 0.99996
[6 empty] | 2.4242 | 1.833e-12 | 0.00010073 | 0.9999
[6 empty] | 2.7475 | 3.8898e-13 | 0.00023116 | 0.99977
[6 empty] | 3.0707 | 7.8865e-14 | 0.00050775 | 0.99949
[6 empty] | 3.3939 | 1.5276e-14 | 0.0010678 | 0.99893
[6 empty] | 3.7172 | 2.8267e-15 | 0.0021504 | 0.99785
[6 empty] | 4.0404 | 4.9969e-16 | 0.0041486 | 0.99585
[6 empty] | 4.3636 | 8.4378e-17 | 0.0076702 | 0.99233
[6 empty] | 4.6869 | 1.361e-17 | 0.013596 | 0.9864
[6 empty] | 5.0101 | 2.097e-18 | 0.023116 | 0.97688
[6 empty] | 5.3333 | 3.0862e-19 | 0.03772 | 0.96228
[6 empty] | 5.6566 | 4.3381e-20 | 0.05911 | 0.94089
[6 empty] | 5.9798 | 5.8244e-21 | 0.089022 | 0.91098
[6 empty] | 6.303 | 7.4688e-22 | 0.12896 | 0.87104
[6 empty] | 6.6263 | 9.1472e-23 | 0.17988 | 0.82012
[6 empty] | 6.9495 | 1.07e-23 | 0.24186 | 0.75814
[6 empty] | 7.2727 | 1.1953e-24 | 0.31389 | 0.68611
[6 empty] | 7.596 | 1.2752e-25 | 0.39383 | 0.60617
[6 empty] | 7.9192 | 1.2993e-26 | 0.47852 | 0.52148
[6 empty] | 8.2424 | 1.2643e-27 | 0.5642 | 0.4358
[6 empty] | 8.5657 | 1.1748e-28 | 0.64695 | 0.35305
[6 empty] | 8.8889 | 1.0425e-29 | 0.72327 | 0.27673
[6 empty] | 9.2121 | 8.8346e-31 | 0.79048 | 0.20952
[6 empty] | 9.5354 | 7.1495e-32 | 0.84698 | 0.15302
[6 empty] | 9.8586 | 5.525e-33 | 0.89234 | 0.10766
[6 empty] | 10.1818 | 4.0772e-34 | 0.9271 | 0.072898
[6 empty] | 10.5051 | 2.8731e-35 | 0.95254 | 0.047456
[6 empty] | 10.8283 | 1.9334e-36 | 0.97032 | 0.02968
[6 empty] | 11.1515 | 1.2423e-37 | 0.98218 | 0.01782
[6 empty] | 11.4747 | 7.6224e-39 | 0.98973 | 0.010266
[6 empty] | 11.798 | 4.4659e-40 | 0.99433 | 0.0056709
[6 empty] | 12.1212 | 2.4984e-41 | 0.997 | 0.0030028
[6 empty] | 12.4444 | 1.3347e-42 | 0.99848 | 0.0015235
[6 empty] | 12.7677 | 6.8079e-44 | 0.99926 | 0.00074032
[6 empty] | 13.0909 | 3.3158e-45 | 0.99966 | 0.00034447
[6 empty] | 13.4141 | 1.542e-46 | 0.99985 | 0.00015344
[6 empty] | 13.7374 | 6.8475e-48 | 0.99993 | 6.5408e-05
[6 empty] | 14.0606 | 2.9034e-49 | 0.99997 | 2.668e-05
[6 empty] | 14.3838 | 1.1754e-50 | 0.99999 | 1.0411e-05
[6 empty] | 14.7071 | 4.5436e-52 | 1 | 3.8858e-06
[6 empty] | 15.0303 | 1.677e-53 | 1 | 1.3871e-06
[6 empty] | 15.3535 | 5.9099e-55 | 1 | 4.7347e-07
[6 empty] | 15.6768 | 1.9886e-56 | 1 | 1.5452e-07
[6 empty] | 16 | 6.3888e-58 | 1 | 4.8213e-08

## Guardband Calculator
Plot PFA as function of guardband width
← Intermediate and plot data in hidden columns here
Global (Average) Risk Calculator | [5 empty] | GBW | PFA | PFR
Nominal/Expected Value | 0 |  |  | Lower Tolerance | -4 | 0 | 2.9553 | 5.5289
Tolerance | 0 | ± | 4 | Upper Tolerance | 4 | 0.040404 | 2.8204 | 5.8344
In-Tolerance Probability | 85% |  |  |  |  | 0.080808 | 2.6889 | 6.1513
Measurement Uncertainty (k=2) ± | 2 |  |  | Range Lower | -4 | 0.12121 | 2.5607 | 6.4798
|  |  |  | Range Upper | 4 | 0.16162 | 2.436 | 6.8201
[6 empty] | 0.20202 | 2.3149 | 7.1723
[6 empty] | 0.24242 | 2.1974 | 7.5366
|  |  |  | Range | 8 | 0.28283 | 2.0837 | 7.913
|  |  |  | Low | -8 | 0.32323 | 1.9737 | 8.3016
|  |  |  | High | 8 | 0.36364 | 1.8674 | 8.7026
[6 empty] | 0.40404 | 1.765 | 9.116
|  |  |  | sigma product | 2.7787 | 0.44444 | 1.6663 | 9.5418
[6 empty] | 0.48485 | 1.5715 | 9.9802
[6 empty] | 0.52525 | 1.4804 | 10.431
[6 empty] | 0.56566 | 1.393 | 10.8944
[6 empty] | 0.60606 | 1.3094 | 11.3704
[6 empty] | 0.64646 | 1.2294 | 11.8588
[6 empty] | 0.68687 | 1.1531 | 12.3598
[6 empty] | 0.72727 | 1.0802 | 12.8733
[6 empty] | 0.76768 | 1.0109 | 13.3991
[6 empty] | 0.80808 | 0.94494 | 13.9373
[6 empty] | 0.84848 | 0.88231 | 14.4879
[6 empty] | 0.88889 | 0.82291 | 15.0507
[6 empty] | 0.92929 | 0.76665 | 15.6257
[6 empty] | 0.9697 | 0.71343 | 16.2127
[6 empty] | 1.0101 | 0.66317 | 16.8117
[6 empty] | 1.0505 | 0.61575 | 17.4227
[6 empty] | 1.0909 | 0.57108 | 18.0454
[6 empty] | 1.1313 | 0.52905 | 18.6799
[6 empty] | 1.1717 | 0.48956 | 19.3259
[6 empty] | 1.2121 | 0.45251 | 19.9834
[6 empty] | 1.2525 | 0.41778 | 20.6522
[6 empty] | 1.2929 | 0.38529 | 21.3323
[6 empty] | 1.3333 | 0.35491 | 22.0235
[6 empty] | 1.3737 | 0.32656 | 22.7257
[6 empty] | 1.4141 | 0.30013 | 23.4387
[6 empty] | 1.4545 | 0.27553 | 24.1625
[6 empty] | 1.4949 | 0.25265 | 24.8969
[6 empty] | 1.5354 | 0.23141 | 25.6417
[6 empty] | 1.5758 | 0.21171 | 26.3969
[6 empty] | 1.6162 | 0.19347 | 27.1623
[6 empty] | 1.6566 | 0.17659 | 27.9377
[6 empty] | 1.697 | 0.16101 | 28.7231
[6 empty] | 1.7374 | 0.14663 | 29.5183
[6 empty] | 1.7778 | 0.13338 | 30.3231
[6 empty] | 1.8182 | 0.12118 | 31.1375
[6 empty] | 1.8586 | 0.10998 | 31.9612
[6 empty] | 1.899 | 0.099693 | 32.7942
[6 empty] | 1.9394 | 0.090266 | 33.6363
[6 empty] | 1.9798 | 0.081635 | 34.4873
[6 empty] | 2.0202 | 0.073745 | 35.3472
[6 empty] | 2.0606 | 0.066539 | 36.2157
[6 empty] | 2.101 | 0.059968 | 37.0928
[6 empty] | 2.1414 | 0.053983 | 37.9782
[6 empty] | 2.1818 | 0.048539 | 38.8719
[6 empty] | 2.2222 | 0.043593 | 39.7737
[6 empty] | 2.2626 | 0.039106 | 40.6835
[6 empty] | 2.303 | 0.035039 | 41.6011
[6 empty] | 2.3434 | 0.031359 | 42.5263
[6 empty] | 2.3838 | 0.028033 | 43.459
[6 empty] | 2.4242 | 0.02503 | 44.3991
[6 empty] | 2.4646 | 0.022322 | 45.3464
[6 empty] | 2.5051 | 0.019884 | 46.3008
[6 empty] | 2.5455 | 0.017692 | 47.262
[6 empty] | 2.5859 | 0.015723 | 48.23
[6 empty] | 2.6263 | 0.013957 | 49.2045
[6 empty] | 2.6667 | 0.012374 | 50.1855
[6 empty] | 2.7071 | 0.010958 | 51.1727
[6 empty] | 2.7475 | 0.0096927 | 52.166
[6 empty] | 2.7879 | 0.0085633 | 53.1652
[6 empty] | 2.8283 | 0.0075565 | 54.1702
[6 empty] | 2.8687 | 0.0066602 | 55.1807
[6 empty] | 2.9091 | 0.0058632 | 56.1967
[6 empty] | 2.9495 | 0.0051554 | 57.2179
[6 empty] | 2.9899 | 0.0045276 | 58.2442
[6 empty] | 3.0303 | 0.0039714 | 59.2753
[6 empty] | 3.0707 | 0.0034793 | 60.3112
[6 empty] | 3.1111 | 0.0030443 | 61.3516
[6 empty] | 3.1515 | 0.0026603 | 62.3963
[6 empty] | 3.1919 | 0.0023217 | 63.4452
[6 empty] | 3.2323 | 0.0020235 | 64.4981
[6 empty] | 3.2727 | 0.0017611 | 65.5547
[6 empty] | 3.3131 | 0.0015303 | 66.615
[6 empty] | 3.3535 | 0.0013277 | 67.6786
[6 empty] | 3.3939 | 0.0011498 | 68.7455
[6 empty] | 3.4343 | 0.00099366 | 69.8154
[6 empty] | 3.4747 | 0.0008567 | 70.8881
[6 empty] | 3.5152 | 0.00073651 | 71.9635
[6 empty] | 3.5556 | 0.00063097 | 73.0413
[6 empty] | 3.596 | 0.00053816 | 74.1213
[6 empty] | 3.6364 | 0.00045636 | 75.2033
[6 empty] | 3.6768 | 0.00038401 | 76.2872
[6 empty] | 3.7172 | 0.00031973 | 77.3728
[6 empty] | 3.7576 | 0.00026223 | 78.4597
[6 empty] | 3.798 | 0.00021035 | 79.5479
[6 empty] | 3.8384 | 0.00016301 | 80.6371
[6 empty] | 3.8788 | 0.00011923 | 81.7271
[6 empty] | 3.9192 | 7.8058e-05 | 82.8177
[6 empty] | 3.9596 | 3.8604e-05 | 83.9088
[6 empty] | 4 | 0 | 85

## Risk Curves
Plot PFA and PFR versus ITP at different TURs, with optional guardbanding
ITP Low | 45
ITP High | 95
TUR Values | 1.5 | 2 | 3 | 4 | 1.5 | 2 | 3 | 4
Guardband | None
Manual Guardband Factor | 0.9
ITP
45 | 4.8823 | 3.7468 | 2.5532 | 1.9352 | 5.6589 | 4.1886 | 2.7513 | 2.0469
45.5051 | 4.9115 | 3.7712 | 2.5713 | 1.9494 | 5.7138 | 4.2279 | 2.776 | 2.0649
46.0101 | 4.9398 | 3.7951 | 2.5889 | 1.9633 | 5.7685 | 4.2668 | 2.8004 | 2.0826
46.5152 | 4.9672 | 3.8183 | 2.6062 | 1.9769 | 5.8228 | 4.3055 | 2.8246 | 2.1002
47.0202 | 4.9937 | 3.8408 | 2.6231 | 1.9903 | 5.8768 | 4.3438 | 2.8486 | 2.1175
47.5253 | 5.0194 | 3.8628 | 2.6395 | 2.0033 | 5.9304 | 4.3818 | 2.8724 | 2.1347
48.0303 | 5.0441 | 3.8841 | 2.6556 | 2.0161 | 5.9837 | 4.4195 | 2.8958 | 2.1517
48.5354 | 5.0679 | 3.9047 | 2.6713 | 2.0286 | 6.0366 | 4.4569 | 2.9191 | 2.1684
49.0404 | 5.0908 | 3.9247 | 2.6865 | 2.0407 | 6.0891 | 4.4939 | 2.942 | 2.1849
49.5455 | 5.1127 | 3.944 | 2.7014 | 2.0526 | 6.1413 | 4.5306 | 2.9647 | 2.2013
50.0505 | 5.1337 | 3.9626 | 2.7158 | 2.0642 | 6.193 | 4.5669 | 2.9871 | 2.2174
50.5556 | 5.1537 | 3.9806 | 2.7298 | 2.0754 | 6.2444 | 4.6029 | 3.0093 | 2.2332
51.0606 | 5.1728 | 3.9978 | 2.7433 | 2.0864 | 6.2954 | 4.6385 | 3.0311 | 2.2489
51.5657 | 5.1908 | 4.0143 | 2.7564 | 2.097 | 6.346 | 4.6738 | 3.0527 | 2.2643
52.0707 | 5.2079 | 4.0301 | 2.769 | 2.1073 | 6.3961 | 4.7087 | 3.074 | 2.2795
52.5758 | 5.224 | 4.0452 | 2.7812 | 2.1172 | 6.4459 | 4.7432 | 3.0949 | 2.2944
53.0808 | 5.239 | 4.0595 | 2.7929 | 2.1269 | 6.4952 | 4.7773 | 3.1156 | 2.3091
53.5859 | 5.253 | 4.0732 | 2.8042 | 2.1362 | 6.544 | 4.811 | 3.1359 | 2.3235
54.0909 | 5.266 | 4.086 | 2.815 | 2.1451 | 6.5924 | 4.8443 | 3.156 | 2.3377
54.596 | 5.2779 | 4.0981 | 2.8253 | 2.1537 | 6.6404 | 4.8772 | 3.1757 | 2.3516
55.101 | 5.2888 | 4.1094 | 2.835 | 2.1619 | 6.6878 | 4.9097 | 3.1951 | 2.3653
55.6061 | 5.2986 | 4.1199 | 2.8443 | 2.1697 | 6.7348 | 4.9417 | 3.2141 | 2.3787
56.1111 | 5.3073 | 4.1297 | 2.8531 | 2.1772 | 6.7813 | 4.9733 | 3.2328 | 2.3918
56.6162 | 5.3149 | 4.1386 | 2.8614 | 2.1843 | 6.8273 | 5.0044 | 3.2512 | 2.4046
57.1212 | 5.3214 | 4.1467 | 2.8692 | 2.1911 | 6.8728 | 5.0351 | 3.2692 | 2.4171
57.6263 | 5.3268 | 4.154 | 2.8764 | 2.1974 | 6.9178 | 5.0653 | 3.2868 | 2.4294
58.1313 | 5.331 | 4.1605 | 2.8831 | 2.2034 | 6.9622 | 5.0951 | 3.304 | 2.4413
58.6364 | 5.3342 | 4.1661 | 2.8892 | 2.2089 | 7.0061 | 5.1243 | 3.3209 | 2.4529
59.1414 | 5.3361 | 4.1708 | 2.8948 | 2.2141 | 7.0494 | 5.153 | 3.3374 | 2.4643
59.6465 | 5.3369 | 4.1747 | 2.8998 | 2.2188 | 7.0922 | 5.1813 | 3.3535 | 2.4753
60.1515 | 5.3365 | 4.1778 | 2.9043 | 2.2232 | 7.1343 | 5.209 | 3.3692 | 2.486
60.6566 | 5.335 | 4.1799 | 2.9082 | 2.2271 | 7.1759 | 5.2362 | 3.3844 | 2.4963
61.1616 | 5.3322 | 4.1811 | 2.9115 | 2.2305 | 7.2169 | 5.2628 | 3.3993 | 2.5064
61.6667 | 5.3282 | 4.1815 | 2.9142 | 2.2335 | 7.2572 | 5.2889 | 3.4137 | 2.516
62.1717 | 5.323 | 4.1809 | 2.9163 | 2.2361 | 7.2969 | 5.3144 | 3.4277 | 2.5253
62.6768 | 5.3165 | 4.1793 | 2.9178 | 2.2383 | 7.336 | 5.3393 | 3.4412 | 2.5343
63.1818 | 5.3088 | 4.1769 | 2.9186 | 2.2399 | 7.3744 | 5.3636 | 3.4543 | 2.5429
63.6869 | 5.2999 | 4.1735 | 2.9189 | 2.2412 | 7.4121 | 5.3873 | 3.4669 | 2.5511
64.1919 | 5.2896 | 4.1691 | 2.9185 | 2.2419 | 7.4491 | 5.4104 | 3.479 | 2.559
64.697 | 5.2781 | 4.1637 | 2.9174 | 2.2421 | 7.4853 | 5.4329 | 3.4906 | 2.5664
65.202 | 5.2653 | 4.1573 | 2.9157 | 2.2419 | 7.5209 | 5.4547 | 3.5018 | 2.5735
65.7071 | 5.2511 | 4.1499 | 2.9133 | 2.2412 | 7.5557 | 5.4758 | 3.5124 | 2.5801
66.2121 | 5.2356 | 4.1415 | 2.9102 | 2.2399 | 7.5897 | 5.4963 | 3.5225 | 2.5863
66.7172 | 5.2188 | 4.1321 | 2.9065 | 2.2382 | 7.6229 | 5.516 | 3.532 | 2.5921
67.2222 | 5.2006 | 4.1216 | 2.902 | 2.2359 | 7.6553 | 5.535 | 3.541 | 2.5975
67.7273 | 5.1811 | 4.1101 | 2.8969 | 2.2331 | 7.6869 | 5.5533 | 3.5494 | 2.6024
68.2323 | 5.1601 | 4.0975 | 2.891 | 2.2297 | 7.7176 | 5.5708 | 3.5573 | 2.6069
68.7374 | 5.1378 | 4.0838 | 2.8844 | 2.2258 | 7.7475 | 5.5876 | 3.5646 | 2.6108
69.2424 | 5.114 | 4.069 | 2.877 | 2.2214 | 7.7764 | 5.6035 | 3.5712 | 2.6143
69.7475 | 5.0888 | 4.0531 | 2.8689 | 2.2163 | 7.8045 | 5.6187 | 3.5773 | 2.6174
70.2525 | 5.0622 | 4.0361 | 2.8599 | 2.2107 | 7.8315 | 5.633 | 3.5827 | 2.6199
70.7576 | 5.0341 | 4.0179 | 2.8503 | 2.2045 | 7.8576 | 5.6464 | 3.5874 | 2.6219
71.2626 | 5.0045 | 3.9985 | 2.8398 | 2.1977 | 7.8827 | 5.659 | 3.5915 | 2.6233
71.7677 | 4.9735 | 3.978 | 2.8285 | 2.1903 | 7.9068 | 5.6706 | 3.5949 | 2.6242
72.2727 | 4.9409 | 3.9563 | 2.8164 | 2.1823 | 7.9298 | 5.6813 | 3.5975 | 2.6246
72.7778 | 4.9068 | 3.9334 | 2.8034 | 2.1736 | 7.9517 | 5.691 | 3.5995 | 2.6244
73.2828 | 4.8712 | 3.9092 | 2.7896 | 2.1643 | 7.9724 | 5.6998 | 3.6007 | 2.6236
73.7879 | 4.834 | 3.8838 | 2.7749 | 2.1543 | 7.992 | 5.7075 | 3.6011 | 2.6222
74.2929 | 4.7952 | 3.8571 | 2.7593 | 2.1436 | 8.0104 | 5.7142 | 3.6007 | 2.6201
74.798 | 4.7549 | 3.8292 | 2.7429 | 2.1323 | 8.0276 | 5.7197 | 3.5996 | 2.6175
75.303 | 4.7129 | 3.7999 | 2.7255 | 2.1202 | 8.0435 | 5.7242 | 3.5976 | 2.6142
75.8081 | 4.6693 | 3.7693 | 2.7072 | 2.1075 | 8.058 | 5.7275 | 3.5947 | 2.6102
76.3131 | 4.6241 | 3.7374 | 2.6879 | 2.094 | 8.0712 | 5.7296 | 3.5909 | 2.6055
76.8182 | 4.5772 | 3.7041 | 2.6677 | 2.0798 | 8.083 | 5.7305 | 3.5862 | 2.6001
77.3232 | 4.5286 | 3.6695 | 2.6464 | 2.0648 | 8.0933 | 5.7301 | 3.5806 | 2.5939
77.8283 | 4.4783 | 3.6334 | 2.6242 | 2.049 | 8.1021 | 5.7284 | 3.574 | 2.587
78.3333 | 4.4263 | 3.5959 | 2.601 | 2.0325 | 8.1094 | 5.7254 | 3.5664 | 2.5793
78.8384 | 4.3725 | 3.557 | 2.5767 | 2.0151 | 8.115 | 5.7209 | 3.5577 | 2.5708
79.3434 | 4.3169 | 3.5166 | 2.5513 | 1.9969 | 8.1189 | 5.715 | 3.548 | 2.5615
79.8485 | 4.2596 | 3.4747 | 2.5248 | 1.9779 | 8.1211 | 5.7075 | 3.5371 | 2.5513
80.3535 | 4.2004 | 3.4313 | 2.4973 | 1.958 | 8.1214 | 5.6985 | 3.5251 | 2.5401
80.8586 | 4.1394 | 3.3863 | 2.4686 | 1.9372 | 8.1199 | 5.6878 | 3.5119 | 2.5281
81.3636 | 4.0766 | 3.3397 | 2.4387 | 1.9155 | 8.1163 | 5.6754 | 3.4974 | 2.5151
81.8687 | 4.0118 | 3.2916 | 2.4076 | 1.8928 | 8.1107 | 5.6613 | 3.4816 | 2.5011
82.3737 | 3.9451 | 3.2418 | 2.3754 | 1.8693 | 8.103 | 5.6453 | 3.4645 | 2.4861
82.8788 | 3.8765 | 3.1903 | 2.3418 | 1.8447 | 8.093 | 5.6273 | 3.446 | 2.4699
83.3838 | 3.8059 | 3.1372 | 2.307 | 1.8191 | 8.0807 | 5.6074 | 3.426 | 2.4527
83.8889 | 3.7333 | 3.0823 | 2.2709 | 1.7925 | 8.0658 | 5.5854 | 3.4045 | 2.4343
84.3939 | 3.6587 | 3.0256 | 2.2335 | 1.7648 | 8.0484 | 5.5611 | 3.3814 | 2.4147
84.899 | 3.582 | 2.9672 | 2.1947 | 1.7361 | 8.0283 | 5.5345 | 3.3567 | 2.3938
85.404 | 3.5032 | 2.9069 | 2.1545 | 1.7062 | 8.0053 | 5.5055 | 3.3302 | 2.3716
85.9091 | 3.4222 | 2.8447 | 2.1128 | 1.6752 | 7.9793 | 5.4739 | 3.3018 | 2.3479
86.4141 | 3.3391 | 2.7806 | 2.0696 | 1.6429 | 7.9501 | 5.4397 | 3.2716 | 2.3229
86.9192 | 3.2538 | 2.7146 | 2.0249 | 1.6094 | 7.9176 | 5.4026 | 3.2393 | 2.2962
87.4242 | 3.1662 | 2.6465 | 1.9786 | 1.5747 | 7.8814 | 5.3624 | 3.2049 | 2.268
87.9293 | 3.0763 | 2.5764 | 1.9307 | 1.5386 | 7.8415 | 5.3191 | 3.1682 | 2.2381
88.4343 | 2.9841 | 2.5041 | 1.8811 | 1.5012 | 7.7975 | 5.2724 | 3.1291 | 2.2064
88.9394 | 2.8895 | 2.4297 | 1.8298 | 1.4623 | 7.7492 | 5.2221 | 3.0876 | 2.1728
89.4444 | 2.7924 | 2.353 | 1.7767 | 1.422 | 7.6962 | 5.1678 | 3.0433 | 2.1371
89.9495 | 2.6928 | 2.2741 | 1.7217 | 1.3801 | 7.6382 | 5.1095 | 2.9961 | 2.0994
90.4545 | 2.5907 | 2.1927 | 1.6648 | 1.3367 | 7.5749 | 5.0467 | 2.9459 | 2.0593
90.9596 | 2.4859 | 2.1089 | 1.6058 | 1.2916 | 7.5057 | 4.9791 | 2.8924 | 2.0168
91.4646 | 2.3784 | 2.0226 | 1.5448 | 1.2447 | 7.4303 | 4.9062 | 2.8353 | 1.9717
91.9697 | 2.2682 | 1.9337 | 1.4816 | 1.196 | 7.3478 | 4.8277 | 2.7743 | 1.9237
92.4747 | 2.155 | 1.842 | 1.4161 | 1.1454 | 7.2578 | 4.743 | 2.7091 | 1.8726
92.9798 | 2.039 | 1.7475 | 1.3481 | 1.0928 | 7.1593 | 4.6514 | 2.6394 | 1.8182
93.4848 | 1.9198 | 1.6501 | 1.2777 | 1.038 | 7.0515 | 4.5522 | 2.5646 | 1.7601
93.9899 | 1.7975 | 1.5496 | 1.2045 | 0.98089 | 6.9331 | 4.4446 | 2.4842 | 1.698
94.4949 | 1.6719 | 1.4458 | 1.1286 | 0.92137 | 6.8028 | 4.3273 | 2.3975 | 1.6314
95 | 1.5428 | 1.3386 | 1.0496 | 0.85924 | 6.6588 | 4.1992 | 2.3038 | 1.5597
← Intermediate and plot data in hidden rows above here

## Cost Model
Cost Model - compare financial tradeoff between false accept and false reject
Global (Average) Risk Calculator |  |  |  | Cost Model |  |  |  | ← Intermediate calculations in hidden columns here.
PFA Required | 0.1% |  |  | Cost of False Accept | 200
Nominal/Expected Value | 0 |  |  | Cost of False Reject | 30 | Lower Tolerance | -1
Tolerance | 0 | ± | 1 | Number of units produced | 1000000 | Upper Tolerance | 1
In-Tolerance Probability | 90% |  |  |  |  | Process Stdev | 0.60796
Measurement Uncertainty (k=2) ± | 0.04
[6 empty] | Guardband Width | 0.01334
Risk Results |  |  |  | Cost Results |  | Lower Acceptance Limit | -0.98666
TUR | 25 |  |  | Number of Falsely Accepted Items | 999.9888 | Upper Acceptance Limit | 0.98666
PFA (No Guardband) | 0.2622% |  |  | Number of Falsely Rejected Items | 5797.4885
PFR (No Guardband) | 0.2807% |  |  | Cost of Falsely Accepted Items | 199997.7619
Acceptance Limit | 0 | ± | 0.98666 | Cost of Falsely Rejected Items | 173924.6559
PFA (Guardbanded) | 0.1% |  |  | Total cost of false decisions | 373922.4178
PFR (Guardbanded) | 0.5797%

## Lookups (hidden)
Tolerance Types | Guardband Methods
± | RDS
to | PFA
max | Manual
min | None